VTiQ	Firmware 	V1	01.12.001
 		V2	01.11.001

Release date: 27 january 2021

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

- In the previous version the version number of the μP2 (v2) firmware version was displayed 
  incorrectly as 01.10.000 instead of 1.11.000. 
  This has been solved. Note that the correct version number is now 01.11.001.

VTiQ	Firmware 	V1	01.12.001
 		V2	01.11.000

Release date: 23 november 2020

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

- Under certain circumstances the viscosity value was not correctly displayed on the
  instruments touchscreen when the viscosity unit was set to P (=Poise) or cP (centiPoise). 
  This has been solved.
  Note: The viscosity value displayed in RheoWin was always correct.

VTiQ	Firmware 	V1	01.12.000
 		V2	01.11.000

Release date: 03 may 2018

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

- Under certain circumstances (sample with high G* value) the reported torque value in 
  OSC-mode was not correctly calculated. This has been solved.
  Note: All other measurement values were correct.

VTiQ	Firmware 	V1	01.12.000
 		V2	01.10.000

Release date: 30 october 2017

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

- The data acquisition for rotational measurements has been improved.

- In manual control mode setting a shear rate or shear stress value selected from a row 
  of values defined with RheoApp did not work correctly. This has been solved.

VTiQ	Firmware 	V1	01.11.000
 		V2	01.09.000

Release date: 12 october 2016

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

- The quality of the data from oscillatory measurements has been improved (improved
  routines for phase angle calculation, compliance correction, etc.)

- The CR-control loop has been improved and now offers a (much) quicker response time for 
  high(er) viscous samples.

- The bearing friction correction now also corrects for changes in the motor temperature.

- The data acqusition during CR and CS rotational ramps has been improved.

- The cone angle value for cone&plate measuring geometries was not displayed correctly.
  This has been solved.

- The temperature of the Peltier heatsink is now continuously monitored. When the heatsink
  temperature reaches an unexpected value (due to insufficient cooling), the Peltier 
  temperature control will be halted and a message will be displayed on the touchscreen.

- The option to automatically activate the Peltier temperature control after the instrument is 
  switched on can now be activate from the touchscreen user interface in 
  the Configuration > Device settings menu.

- It was not possible to enter a geometry factor value higher than 1e+6 using the touchscreen 
  user interface. This has been solved.

VTiQ	Firmware 	V1	01.10.000
 		V2	01.07.000

Release date: 10 may 2016

Below is a list of things that are new and/or changed in this version compared with 
the previous version.
- On the Login screen the default user was not disabled when other users are defined 
 (using the RheoApp software). This has been solved.
 Please note: This version must be used with RheoApp version 2.4.2.9 or higher.

VTiQ	Firmware 	V1	01.09.000
 		V2	01.07.000

Release date: 27 november 2015

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

- There was a bug in firmware V2 01.06.000 regarding the torque mapping.
  This has been solved.

Please note: When updating the firmware files, the HMI update must be applied as well.

VTiQ	Firmware 	V1	01.09.000
 		V2	01.06.000

Release date: 20 november 2015

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

- Operating the VTiQ without using Connect Assist, that is the automatic geometry connection, is now possible by activating the Configuration > Device settings > Allow measurements without geometry recognition checkbox.

- The gap setting routine for cone and plate as well parallel plate geometries has been modified. Please note that the gap setting tool wiht lever arm is needed for this to work as intended.

- The measurement routine for the new Fill Assist tool is now implemented.

- The CR-mode control loop has been optimized for use with high viscous samples.

- The firmware is now compatible with the new pressure cell stand and pressure cell geometries etc.

- The functionality of the oscillation mode has been improved.

- Under certain circumstances the data of an oscillation frequency sweep measurement were not correct for f > 1,5 Hz. This has been solved.

VTiQ	Firmware 	V1	01.07.000
 		V2	01.04.000

Release date: 27 april 2015

- Several improvements in the CD-Oscillation mode.

- Bugfix for an internal communication problem.

VTiQ	Firmware 	V1	01.06.000
 		V2	01.03.000

Release date: 16 march 2015

- New language for HMI: Indonesian.

----------

Contact Thermo Fisher Scientific

If you have any questions and or suggestions regarding the MARS (firmware)

please send an e-mail to:

support.mc.de@thermofisher.com

----------

This file was last updated on: 27.01.2021 by Jint Nijman.