HAAKE RheoWin Version 2.97

Contents of this document

1 Downloading and installing

1.1 Download the RheoWin files

1.2 Installation of RheoWin

1.3 Download the Demo KEY-Diskette

2 What's new

2.1 Version 2.9

2.2 Version 2.91

2.3 Version 2.92

2.4 Version 2.93

2.5 Version 2.94

2.6 Version 2.95

2.7 Version 2.96

2.8 Version 2.97

3 Limitations and known problems

4 Contact Thermo Electron

============================================================

1 Downloading and installing

This is a fully executable version of RheoWin. Please note that in order to be able to install

RheoWin on your computer you need a RheoWin KEY-Diskette !

RheoWin currently supports the following HAAKE rheometers and viscometers:

RS600, RS300, RS150, RS80, RS75, RS100-1Ncm, RS100-5Ncm, RS50, RT20, RT10, VT550,

VT500, VT501, VT6L, VT6R, VT7L, VT7R, RV1, RS1, RW1 fluids, RW1 melts, RV20+RC20/RV30 
with CV20, CV100, M5 or M10.

RheoWin runs under Windows 95, Windows 98, Windows 98 SE, Windows ME, Windows NT4.0,

Windows 2000 and Windows XP; it will NOT run under Windows 3.x or Windows NT3.x.

HAAKE highly recommends the use of a PC with a Pentium II 300 MHz (or faster) processor

and at least 64 Mb of RAM.

Existing users

HAAKE customers who already have a copy of RheoWin should use the RheoWin

Key-Diskette that came with their original copy of RheoWin.

(The KEY-Diskette is independent of the software version).

New users

With the KEY-Diskette that you can download (see below) from this page you can install a

demo-version of RheoWin. The only limitation of this demo-version is that you can not run

measurements with real rheometers/viscometers.

1.1 Download the RheoWin files

Download all eight of the following *.ZIP files and copy them to a single (temporary) directory

on your hard drive.

Files : Rhw297_1.zip, Rhw297_2.zip, Rhw297_3.zip, Rhw297_4.zip, Rhw297_5.zip,

Rhw297_6.zip, Rhw297_7.zip, Rhw297_8.zip.

Then unzip each of the eight *.ZIP files using any suitable unzip program. In case you do not

have an (UN)ZIP program you can download a freeware ZIP program called FreeZip from the

following Web-site : "http://www.ozemail.com.au/~nulifetv/freezip" .

IMPORTANT: Make sure that the unzip program restores the following original subdirectories

\setupdir\0007,   \setupdir\0009, \setupdir\040c,  \setupdir\0013 and  \setupdir\000a

contained in the file Rhw297_1.zip !!

Because of the size of the *.cab files it is not possible anymore to copy the installation files to floppy disks.

1.2 Installation of RheoWin

Although it is not necessary to deinstall an older version (of RheoWin) when you are upgrading

RheoWin to the newest version we do recommend that you completely deinstall the older

version. We also recommend that you reboot the computer after deinstalling the older version

and before installing the new version.

It is necessary to reboot your computer (only if you ran RheoWin after the computer was booted

the last time) before you run the RheoWin installation program !

Here's how to install RheoWin:

From a directory on your harddisk in which you unzipped the eight *.ZIP files.

To install RheoWin from harddisk go to the (temporary) directory in which

you unzipped the eight *.ZIP files after downloading and start the Setup.exe program.

The installation program will guide you through the installation. Please make sure that you

have a KEY-Diskette (see above) readily available since the installation program

will ask for this disk.

1.3 Download the Demo KEY-Diskette

In order to be able to install RheoWin on your computer you will need a RheoWin KEY-Diskette.

You can download a KEY-Diskette for a Demo-version of RheoWin from the link below. After

downloading the file register.zip unzip the file, this will result in a file Register.exe. Then copy

the file Register.exe to a floppy disk and label the disk "RheoWin KEY-Diskette".

Please note:  If you already have a valid RheoWin KEY-Diskette there is no need

to download this file.

----------

2 What's new

2.1 Version 2.9

Below is a list of things that are new compared with version 2.84 of RheoWin.

- General

- RheoWin can now directly read data files created with the HAAKE (DOS) ROT software 
    versions 2.0 to 3.3.

- After clicking the new Information button in the Info about window (Help menu), 
    RheoWin will display a list off all files that are currently stored in the RheoWin directory 
    and all of its sub-directories. The list can be saved in a file.

- Rheometer drivers

- RT10
    - There were problems with oscillation elements in combination with Remote 
      controlled temperature controllers. This has been solved.

- RV1

- There is now a delay of 2 seconds (in a measurement / job) after the torque offset 
      correction has been determined.

(See the option Torque offset correction in the Options tab of the DeviceManager).

- Some problems with printing from "internal" jobs (Data dump element) have been solved.

- RS1, RV1, RW1, RS300
    - The values of the instrument limits (torque range, etc) that are stored in the driver DLL file
      have been corrected.

- There were problems with setting temperature offsets and uploading data to the instrument. 
      This has been solved.

- RS100, RS50

- There were problems with the oscillation mode in the version 2.82 to 2.84. 
      This has been solved.
    - There were problems when reading negative temperature values from analog controlled
      circulators (e.g. F3C). This has been solved.
    - There were timing problems when switching from a ROT measuring element to an 
      OSC measuring element. This has been solved.

- RS300

- The lift had problems moving into measurement position with high viscous samples. 
      This has been solved.

- The Upload graphic definition element now works with the RS300.

- TP1800

- The driver is now compatible with the newest version of the TP1800 control-box 
      (new Eurotherm controller).
  - Analog circulators (F3C)
    - There was a bug in the temperature read routines. This has been solved.

- DC30, DC50
    - The set values for the highest and lowest allowed temperature are now used in RheoWin
      in order to overcome a bug in the circulator firmware.
  - Phoenix

- There was a problem when using a Phoenix circulator connected Remote to a RV1, 
      RS1 or RW1.
      This has been solved.

- TC81, TC501

- The temperature read offset values (tables) have been removed.

- TCP/P
    - The lowest allowed temperature is now - 42°C.

- DeviceManager

- MSC calibration
    - The graph displayed at the end of the test measurement will now show the same
      information as that at the end of the calibration measurement.

- Auxillary devices

- A driver for a special humidity measuring instrument is now part of RheoWin (but not 
      automatically installed).
    - The driver for the humidity measuring instrument can now be used in multiple processes
      at the same time.

- DeviceManager

- RheoWin logon
    - The User name can now be selected from a drop-down list.

- JobManager

- Save data file Element / Filename button

- A problem with saving the data of the last element only, when a fixed file name was 
      selected, has been solved.
    - The new option No data file allows the user not to save measured data.

- Set-Element

- Using this element the motor driveshaft can be positioned at a certain absolute angular 
      position.
    - The angular speed with which the driveshaft is positioned is adjustable.
    - A maximum waiting time for this element can be set.
    - The set-element is listed in the list of element in the protocoll.

- All Time-Elements

- The Trigger functionality (on the Extern tab page) is now only visible and useable when the 
      selected instrument support this functionality.

- Temperature setting : Take from previous

- Even when no temperature controller is selected the correct set temperature values

are displayed in the element editors.

- Message Element

- The message element can now be set to close automatically after a certain amount of time.

- Operator name

- The RheoWin logon name is entered automatically for the operator name when the option 
      Logon name on the Operator tab of the Identification element is used.

- Execute window
    - A temperature is now only displayed when a valid temperature signal is available.

- DataManager

- Marker

- There were problems plotting markers when the X-axis contained negatives values

and an Y-axis was scaled logaritmically. This has been solved.

- Log-scaling

- When 0.01 was entered for the minimum of an axis, the axis was actually scaled starting
      from 0.001. This has been solved.

- Tables

- The font for plain text in tables is now set correctly.

- Page layout from data file

- On the General tab of the Preferences dialog (Configuration menu) there is a new option 
      Layout from file when loading. When this option is active, the page layout settings that are
      saved in a RheoWin data file (*.rwd file), which were set in the JobEditor, are automatically 
      used for the new page that is created in the DataManager when the data file is opened.

2.2 Version 2.91

Below is a list of things that are new compared with version 2.9 of RheoWin.

- General

- The "Please reboot the computer" message window is now "always on top".

- The directory select dialog has been changed. This dialog is now based on the standard

windows file open dialog.

- Rheometer drivers

- All instruments

The basic data acquisition routines of all instruments have been optimized. As a 
    consequence the data acquisition is now faster (more data points per second) and 
    the response time of the RheoWin user interface has improved.

- RS300

Some chances have been made for the Normal force option.

- DeviceManager

- MSC calibration for a rheometer

When the MSC button in the DeviceManager is clicked while the option Use MicroStressControl
     is not active (for that instrument), the user is asked if this option should be activated. The MSC
     dialog will only open when this question is answered with Yes, the option will stay active 
     after the MSC dialog has been closed.

- MSC calibration for a sensor

When the MSC button in the DeviceManager is clicked while the option Use MicroStressControl
    is not active (for that instrument), the user is asked if this option should be activated. The MSC 
    dialog will only open when this question is answered with Yes, the option will stay active after 
    the MSC dialog has been closed.

At the end of the calibration the user is asked if the MSC values should be stored in the sensor 
    data. If MSC is not set active for that sensor the user is asked if it should be set active

- MSC calibration
    At the end of MSC calibration measurement the calculated torque offset values are saved in an
    ASCII file with the extension *.msc, where * stand for the instrument or sensor name. 
    If this file already existst it will be overwritten.

- MSC table import/export 
    On the MSC page of the Properties of any sensor, the MSC table can now be imported from 
    and exported to an ASCII file.

- JobManager

- Start device Job
    On the RS300 running a Job in (the firmware of) the device (rheometer), Start device Job, did 
    not work when the option Normal force on the Additional data page of the Properties of RS300
    dialog was activated. This has been solved.
  - JobEditor : Save data file dialog
    The functionality of the Automatic filename option has been extended. RheoWin now allows
    the automatic "composition" of the filename and subdirectory from items like the substance
    name, sample no., date, time, an automatic counter, etc.

- JobEditor : Saving Jobs

When a Job is saved after running it, the data generated during that run is no longer saved in 
    the Job file, this means that existing Job files will no longer increase in size when they are
    saved again.

- Regression-Element

The calculated Tau-OICCC value was not displayed in a Report printout generated by a Job.
    This has been solved.

- DataManager

- Reports
    A seperate Report viewer for saved reports (*.LL files) is now part of RheoWin. It can be 
    started from the JobManager and DataManager Program menu.

2.3 Version 2.92

Below is a list of things that are new compared with version 2.91 of RheoWin.

- Installation program

- Images of the different rheometers are now displayed in the menus during installation.

- The selection of the different rheometers has been modified.
    (Only for those who use the general key, i.e. ThermoHaake employees and general agents)

- RheoWin is now packaged on 7 or 8 "disks". Each "disk" contains upto 4 Mb, therefore it 
    is not possible any more to distribute RheoWin on 1.44 Mb floppy disks !

- Nederlands (Dutch) is available (again) as a language for the installation program.

- General

- Units

- Quantities for the actual gap (h) and a dummy quantity (img) for indexing images have 
      been added. In a data table the quantity img having the value 1 means that this datapoint 
      has a corresponding image.

- The symbol for the first normal stress difference has been corrected („1 instead of „).

- Rheometer drivers

- RheoScope 1

- The last settings for the light intensity, contrast and integration time are now 
      stored when the communication with the instrument is closed. When the 
      communication is restarted these settings are automatically restored.

- RheoSress 600

- This version includes a driver for the new Thermo Haake RheoStress 600 rheometer.

- DeviceManager

- RheoStress 300 and 600

- The option Normal force on the Options tab in the RS300 device editor has been removed. 
      This option is now activated automatically when the RS300 is equipped with the NF-Option.

- The text on the MSC button is now depending on the type of rheometer. For a RS600 with 
      Nf-option the text is now MSCplus + MFC for example. MSCplus is the new MSC with 720

correction points (instead of 100)  per revolution, MFC is the new Micro Force Correction.

- For a RS300 and RS600 with normal force option the determined MFC correction curve

is now displayed in the graph of the MSCplus + MFC calibration (and test) graph.

- Inertia determination

- The inertia determination routine has been modified, it now automatically adapts its settings 
      to "heavy" measuring geometries like the PZ36, PZ38 and PZ39.

- TC81

- The temperatur offset table was not correctly transfered to the control box when

the TC81 is connected directly (Remote) to a RS300 or RS600. This has been solved.
  - TC501

- The temperatur offset table for the upper heater was not correctly transfered to the

control box when the TC501 is connected directly (Remote) to a RS300 or RS600.

This has been solved.

- JobManager

- Manual control window

- For RS300 and RS600 with the normal force option there is a new Nf-reset button.

- The Nf value is now displayed as number below the graph (instead of the time).

- Lift Element

- The maximum values for AutoTension and Normalforce are read from the instruments.

(RS150 : 25 N, RS300/600 : 50 N)

- Reset Element

- There is a new option for re-setting the normal-force signal to zero.

- Camera control Element

- The buttons Run and Service have been renamed in Operation and Maintenance.

- Internal jobs

- The quantities N1 and Y were not stored in *.rwd data files that were measured using the

"Internal Job" mode. This has been solved.

- All ROT Measurement Elements

- The settings for saving images were not copied correctly when copying an element. 
      This has been solved.

- On the Acquisition page in the Element Editor the items in the Save pictures box have 
      been rearranged.

- ROT stationary flow curve and ROT stationary temperature Elements

- Optionally the normal force signal can be displayed in the second on-line graph. This is 
      enabled by selecting the Show N1 option on the Acquisition tab of the element editor.

- OSC and ROT Temperature Ramp Elements

- The allowed values for deltaT/t are now limited to a technical possible range.

- OSC Time Element with Superimposed Stress

- Minor changes in the Element Editor to incrrease the readability of some dialog items.

- JobExecute window

- After clicking on the >>> button the Execute window will now expand and show two tables, 
      with the current pre-set and measured values respectively.
      The values shown depend on the type of element that is running.

- All Evaluation Elements

- User defined QC-result text were not saved and and not displayed in the report window. 
      This has been solved.

- Interpolation Element

- The symbol selected on the QC tab is now saved in a page (*.rwp) file.

- Curve discussion Element

- The colours selected on the layout tab of the element editor were 
      not saved correctly.
      This has been solved.

- DataManager

- Video display window

- The size measurement tool now also displays the measured size in the statusbar 
      of the video window.

- An optional scale is now displayed in the lower right corner of the video-image. 
      This option can be switched on/off in the right mouse button popup menu.

- A single image can now be saved (TIFF format) using the right mouse button popup menu.

- The following functionality has been added to the right mouse button popup menu :

- the video can be split into and saved as individual images (TIFF or BMP format).

- A single image can now be saved using the windows bitmap format also.

- An image can now be deleted from the data file using the Delete item.

- All images can now be saved in one separate *.avi file using the Save as AVI item

- For "playing" the video (using the uu button) the user can now choose between a default frame-rate

of 15 fps and the actual frame-rate at which the images were captured during the measurement.

- Reports

- The serial number of a rheometer is now available in (printed) reports. This field has 
     been added to Standardreport.lst and Measurereport.lst report template files. 
     This field (VariableSerialNumberMeasureDevice) can be added to any report template by 
     using the report template editor.

- Using the various new report templates (List_*.lst) the protocoll information of multiple files can

now be displayed in one report. 
     On the new Report tab of the Preferences dialog (Configuration menu) the user can 
     select which parts of the protocoll will be displayed in a report.
     (Displaying tables of multiple files in one report is not possible)

- Special report templates for displaying the results of selected evaluations only are

now available.
     The fields used in these templates can of course be used in other templates also. 
     This also works for multiple files in a single report.

- Some of the report templates that come with RheoWin have been renamed.

2.4 Version 2.93

Below is a list of things that are new compared with version 2.92 of RheoWin.

This version is mainly for RheoStress 300, RheoStress 1, RheoScope 1, RheoWave 1, 
and RotoVisco 1 on which the firmware has been updated to (file) version 
Rxx_B_02_03_002_dsp_B_02_00_000.h86.

Please note that the above firmware update will delete the MSC table that is stored in the instrument itself.

This table has to be restored using the MSC calibration routine for the instrument (not for a sensor !!)

- General

- Units

- The density was displayed without its unit in the Identification menu as well in the Protocoll and in Reports.
      This has been solved.

- Language file

- A few text strings were missing, they have been added.

- Username
    - The username can now also contain the empty space character.

- JobManager

- All ROT Measurement Elements

- The equation for the centrifugal force correction (for the normal force) was not
      implemented correctly.
      This has been solved.

- When the density is set to zero in the Identification menu the centrifugal force 
      correction is disabled !

(the density is no longer set 1000 kg/m^3 by default when the entered value is 0 !!)

- Rheometer drivers

- RheoStress 1, RheoScope 1

- The determined MSC table is now also stored in a *.msc file (in the RheoWin 
      drivers directory).

- The MSC table now can have 720 entries instead of 100 (MSCplus)

- The lowest torque value that can be entered in all edit fields is now 0.5 mNm.

- The CD <---> AS switch frequency is now 2 Hz

- RotoVisco 1, RheoStress 1, RheoScope 1, RheoWave 1

- The temperature offset table for integrated temperature controllers is now send correctly.

- RheoStress 300

- There were some problems with instruments with the previous released firmware
       version Rxx_B_02_01_011_dsp_B_01_02_045.h86. This has been solved.

- RheoStress 300, RheoStress 600

-  The MFC table is now send correctly.

- The CD <---> AS switch frequency is now 2 Hz

- All Rheometers / Viscometers

In order to overcome communication problems on certain PCs the user can now set a delay
    time (in the order of milliseconds) for the communication protocoll. This setting has to be

made in the drivers.ini file that is located in the RheoWin Drivers directory. The drivers.ini file 
    that comes with RheoWin has a few example entries. An entry for a certain instrument 
    consists of two lines, the first has to contain the name of the instrument drivers DLL
    filename (without the extension .dll), the second line should read ComWait=xx

were xx has to be replaced by a numeric value between 0 and 99, this value represents 
    the delay time in milliseconds. For the ComWait command to be active the semicolon at 
    the beginning of the line has to be removed !!

- DeviceManager

- RotoVisco 1, RheoStress 1, RheoWave1, RheoScope 1, RheoStress 300, RheoStress 600

- On the Firmware tab of the Properties window of these instruments, there is a 
      new Reset button .

- RheoScope 1

- On the Additional data tab of the Properties window the user can now select the 
      Magnification factor of the microscope objective.

- Phoenix, F/N 6/8

- On the General tab of the Properties window of these circulators the user can now select 
      the Baudrate for the serial communication.

- DataManager

- Video display window

- The size measurement tool and the 50 mm scale are displayed in blue for besser visibility.

2.5 Version 2.94

Below is a list of things that are new compared with version 2.93 of RheoWin.

- Installation program

- When the installation program finds a previously installed version of RheoWin that was installed

using the DEMO key, the installation will stop and prompt the user to deinstall the demo version first.

- General

- Quantities/Units

- The actual gap h is now available as a data-channel (mainly for the Axial-Elements).

- Language file

- Polish (Polski) is now available as one of the dialog languages for RheoWin.

- Data display subsystem

- Some internal changes have been made in the data display subsystem (visapi.dll). As a consequence 
      this subsystem is now faster and will not cause the windows user interface to stop working when it

(the visapi.dll) crashes under Windows 2000 and Windows XP.

- The links to web-pages in the JobManager and DataManager help menus have been updated.

- Rheometer and Thermocontroller drivers

- All Rheometers / Viscometers

- In order to overcome communication problems that caused "character overrun" errors on 
      certain PCs, the basic communication routines have been modified. "Character overrun" 
      errors (caused by the PC-hardware) will now no longer lead to the interruption 
      and premature stopping of jobs.

- TCE / Z

- A driver for the new electrical heated cylinder geometry (for the series 1 instruments) is now part of the installation.

- Barcode Job-Start module (bcwacker.exe)

- The dialog language setting of RheoWin is now valid for this module also.

- JobManager

- Lift Element

- With the Series 1 and RheoStress 300/600 rheometers there was a problem with

moving to the Standby position. This has been solved.

- The Start measurement when reached option of the AutoTension function did not work

with the RS300/600. This has been solved.

- CS/CR Rotation Step and CS/CR Rotation Temperature Step Elements

- With the RheoScope 1 the data in the lower graph (showing the measured quantity

reaching its equilibrium value) was not displayed correctly.
      This has been solved.

- Axial Measurement Elements

- There are two new measurement elements for the RheoSress 300 and RheoStress 600 :

Axial Ramp Element

Using this element the lift can be moved from one (user defined) position to another in a

certain (user defined) time. This can be used for simple squeeze flow experiments, by
      measuring the normal force or for measuring the properties of the sample under rotation at
      different levels of immersion of the sensor in the sample by applying a (constant) rotational 
      speed or torque while the lift is moving.

Axial Time Curve Element

Using this element the lift can be moved with a constant (user defined) normal force acting
      on the sample.

In the near futures the Axial Ramp element will be available (with limited functionality) for 
      the series 1 instruments also.

Please note : for these elements to work the instrument has to be equipped with 
                            the following (or newer)

firmware versions  : 	V1:02.05.006-Aug 13 2002

V2:02.00.000-Feb 22 2002

- All Measurement Elements

- RheoWin now saves the actual measured gap for each data point

- In the case of a parallel plate geometry the shear-rate and strain are now calculated

using the actual measured gap for each data point.

Note : This means that with a parallel plate geometry you can now (for example) measure 
                a flow curve, decrease the gap (using a lift element) and measure a second flow 
                curve (with the new gap) all in one Job !

- The two axial elements always use the set gap (??)

- Manual mode

- During manual mode the last set gap is used to calculate rheological quantities (??)

- MultiWave Elements

When no measured temperature was available the software showed an arbitrary value. 
    This has been changed, the software will now show no value.

- Camera control Element

- The camera control window is now also accessible from the ("always on top") JobControl
       window. This allows the operator to modify the camera settings (contrast, brightness
       etc.) during a Job run..

- When the camera control window is used in manual mode (i.e. not as part of Job)
      an oscillation in

CD-mode can be started, images and data are not saved in this operation mode.

- The set values for contrast, brightness, etc are now also displayed as numeric values.

- The automatic integration time setting has not yet been implemented !!

- All camera settings are now accessible even when the objective is not in 
      meaurement position.

- Oscillation measurement Elements

- RheoWin can now save the images acquired during oscillation measurements. 
      In contrast to rotation measurements where exactly one image is acquired and saved
      for each rheological data point, with oscillation measurements more then one image
      per data point can be saved, this depending on the duration of one oscillation period
      and the number of repitions for each data point.

- A "character overrun error" that appeared under certain circumstances and caused problems

is now handled properly.

- With the RheoScope 1 the oscillation raw data (sine-waves) was not displayed correctly.
      This has been solved.

- On the RS300/600 with normal force option the mean value (over one oscillation period) of

the normal force signal is calculated and recorded in the data file. Please note that this

normal force value is not a "correct" rheometric value.

- DataManager

- Video display window

- RheoWin would crash when a file measured with the RheoScope did not contain any

images, for example when all images were deleted manually. This has been solved.

- Reports

- The four Multifile_xxx.lst report template files that come with RheoWin did not work. 
    This has been solved.

2.6 Version 2.95

Below is a list of things that are new compared with version 2.94 of RheoWin.

- Installation program

- The option to add icons for the RheoWin JobManager and DataManager 
    to the Windows TaskBar (QuickLaunch area) is now available in Windows XP also.

- The readme file can now be viewed (using the write.exe program) at the end of the installation

on Windows NT, 2000 and XP also.

- The selection of thermocontrollers and circulators for the RS/RT rheometers has been

separated into two different dialogs.

- General

- Service and maintenance messages

- Service and maintenance messages are now automatically generated. A service request

form can be printed.

- Problem with calculating shear rate

- On instruments without a lift (i.e. M5, M10, VT500, VT501, VT550) and sensors with a

variable gap (Plate/Plate geometry) the shear rate was calculated incorrectly under certain

circumstances.

- Rheometer and Thermocontroller drivers

- UTC driver

- A driver for the new Universal Temperature Controller (UTC) has been added.

- RheoStress 300 and 600 driver

- The temperature signal from the extra PT100 sensor was not measured correctly.

This has been solved.

- RheoStress 1 (RS1) and RheoScope 1 (RO1) drivers

- The lowest torque in oscillation mode is now 0.5 10^-6 Nm.

- Phoenix driver

- There was a problem reading the temperature in combination with the RS600 (Remote

connection). This has been solved.

- DeviceManager

- RheoScope 1 (RO1)

A new image type (for the images contained in the RheoWin *.rwd files) can be selected on

the Additional data page of the Properties of RO1 dialog. 
    The LuraWave image format combines a high image compression level with a very good 
    image quality.

For the influence of the image format and compression rate on the image size, image quality
    and frame rate see the table below :

Image format      Compression rate    Image size (kb)   Image quality           Frame rate (f/s)

TIFFor BMP               n.a.                      300                  lossless                up to 15

LuraWave                  1                        190                  lossless                  3 - 4

LuraWave                 10                         30                  very good                4 - 5

LuraWave                100                          3            visible loss of quality      5 - 6

- Temperature offset tables

- The temperature offset tables in the instruments memory (firmware) are not deleted anymore

when there is no new table available.

- Temperature read offset tables are now available (again) for the DC5, DC30, DC50 and
      FN6/8 circulators as well as the TC80 temperature controller.

- Phoenix circulator

- The switch for external temperature control has been removed. The external control is now

automatically activated when the temperature sensor is connected.

- Automatic backup

- An automatic backup of all the settings in the DeviceManager (which are stored in the

Drivers.flp file in the Rheowin drivers directory) is stored in the Drivers_Autosave.flp file in

the (user-defined) AutoSave directory.

- The settings stored in the backup file can be restored using the new Automatic backup

command from the Restore device definitions menu in the Configurations menu.

- Manual backup

- Using the new Import, Export and Replace buttons the user to create and restore (backup)

copies of the drivers.flp file (RheoWin Drivers directory) which contains all the information

handled by the DeviceManager. The Export command will copy all the settings in a file.

The Replace command will load a file and replace the current settings with the setting

stored in the file. The Import command will load a file and add the settings from that file

to the current settings without overwriting any settings.

- The settings stored in a file can also be loaded by using the new From file command

from the Restore device definitions menu in the Configurations menu. This will overwrite

the current settings.

- JobManager

- JobExecute window
    - The unit for the set deformation (CD-Mode) is now displayed correctly.

- Oscillation Frequency Sweep Element

- The user can now select if the distribution of data points per decade is based on

decades in f [Hz] or decades in † [rad/s].

- Oscillation Elements

- Autotension is now available in the following Elements : OSC-Timecurve, OSC Temperature-
      Ramp, OSC-Temperature-Step, OSC-MultiWave-Temperature-Ramp, OSC-MultiWave-
      Temperature-Step and OSC-Timecurve with superimposed constant stress.

The Autotension functionality is available on new Axial tab in the JobEditor. When 
      Autotension is activated the instrument will adjust the gap in such a way that the normal 
      force acting on the sample is constant throughout the measurement. The normal force

(gap) is only adjusted during a  short pause between two oscillation periods.
      AutoTension has a working range from - 1.0 to 30 N.

- The oscillation frequency range for the Series 1 rheometer, the RheoStress 300 and 600,

has been extended by one decade at the low frequency side.

-  Axial Ramp Element

- This element now also works for increasing gaps. The Axial Elements work with following 
      instruments RotoVisco 1, RheoStress 1, RheoScope 1, RS300 (all versions), 
      RS600 (all versions).

- Camera control Element

- In the camera control window the automatic integration time option is now functional.

- Saving Images

- In the Element-Editors the user can now select whether he wants to switch-on the image

saving automatically (using the parameters in the Element-Editor) or manually by using a new

button in the Job Execute window.

- Save data file Element / Filename button

- The counter of the automatic filename option was not increased when the complete filename

was longer then 128 characters. This has been solved.

- CS/CR Rotation Step and CS/CR Rotation Time Curve Elements

- In the Element-Editor there is a new option available to stop the rheometer motor

automatically each time an image is being taken with the camera (RheoScope 1 only).

- MultiWave Elements

- The temperature was not measured correctly under certain circumstances.

This has been solved.

- Graph during measurement

- When the graph layout dialog was opened (and closed again) during a measurement, in 
      order to change/add axis or curves for example, the graph would only show the data of 
      the last (current) segment for the new axis/curves. This has been changed, the graph will
      now show the data for all segments.

- DataManager

- Video display window

- Single images can now also be saved in the new LuraWave (*.lwf) format using the 
      right mouse button popup menu, when this format was used during the measurement 
      (see DeviceManager).

- RheoWin DataManager crashed when the video display window was closed while a video

was still runing. This has been solved.

- Tables

- There was a problem displaying numbers with 10 significant digits correctly. 
      This has been solved.

2.7 Version 2.96

Version 2.96 fixes a few problems of version 2.95, there are no new features.

- DeviceManager

- All circulator

- When a circulator was manually added (using the Add button in DeviceManager) the interface

could not be selected. This has been solved.

- DataManager

- Saving Images

- Images saved (during a Job) in the compressed LWF format could not be individually saved in

in other formats. This has been solved.

2.8 Version 2.97

Below is a list of things that are new compared with version 2.96 of RheoWin.

- Rheometer and Thermocontroller drivers

- RS75, RS80, RS150

- Character-Overrun errors that occur during an oscillation measurement were not treated 
       correctly. This has been solved.

- RS300, RS600

- The trigger functionality now also works in upload mode (a new firmware is

needed for this work).

- RS1, RheoScope 1, RV1, RW1, RS300, RS600

- Mainly in the Temperatur-Set element (but also in other elements) there was a problem with
      "negative" start times under certain circumstances. This has been solved.

- The service message will now appear again after 30 days.

- All RheoStress instruments, RT10, RT20, RotoVisco 1, RheoScope 1

- The second temperature channel randomly contained "strange" values. This has been solved.

- F/N 6/8 Circulators drivers

- When the "take from previous" option in measurement elements was used, there was a 
      problem with the temperature offset values (the temperature was increased from element
      to element). This has been solved.

- TCE/C

- The alarm-temperature value that is send to the controller was calculated (and send)
      incorrectly. This has been solved.

- TCE/St

- The alarm-temperature value that is send to the controller was calculated (and send)
      incorrectly. This has been solved.
    - The temperature range is now limited to 0 - 100 °C.

- UTC

- The alarm-temperature value that is send to the controller was calculated (and send)
      incorrectly. This has been solved.
    - When a Peltier unit is connected to the UTC box, the second temperatur channel contains 
      the temperature of the heat-exchanger.

- Jobs that are run in Upload-mode now work ok with the UTC.

- DeviceManager

- Sensors

- For the RheoScope 1, the following sensor have been added : C35/1 Ti Po, C60/1 Ti Po,

PP35 Ti Po, PP60 Ti Po. The C70/1 Ti has been renamed to C70/1 Ti Po. "Po" indicates the

polished surface.

- Two parallel plate sensors for electro-rheology (PP35 ER and PP60 ER) have been added.

- Three sensors for the pressure cell (PZ35, PZ37, PZ38 b) have been added.

- There was a problem displaying the properties of ISO2555 sensors. This has been solved

- There was a problem when copying sensor in the DeviceManager. This has been solved.

- JobManager

- Save Element
    - In the Save data file dialog there is a new button "Setting" which adds new possiblities
      to the "Ask for file name" functionality : The user can now define a fixed text block that
      is automatically added to each filename that is entered. This text block can be added as
      a prefix or a suffix.

- Camera control Element

- The automatic integration time parameters were not correctly loaded from a Job-file. 
      This has been corrected.
    - RheoWin crashed when this element was active at the end of a job. This has been solved.

- Lift Element

- On the Rotation page of the Lift-Element editor the Opening gap (lift) option is now available 
      for the RS300 and RS600. With this option the sensor can be set to rotate at a certain 
      speed when the lift is moving downwards (i.e. opening the gap).

- The Include lift element time in meas time option now also works when a Job contains 
      more then one Lift element.

- Rotation Ramp Element

- The data-acquisition at the end of a ramp has been improved.

- Rotation Temperature Ramp and Oscillation Temperature Ramp Element

- The displayed temperature gradient was not calculated correctly. This has been solved.

- Rotation Stationary FlowCurve Element
    - There was problem caused by the display of the normal force signal. This has been solved.

- Reference Element
    - Some display problems in the editor of this element have been solved.

- All Evaluation Elements
    - The segment selection was not working properly with long segment lists. 
      This has been solved.

- DataManager

- Graph dialog
    - Under Windows 2000 and XP there were problems with the selection and display of
      the X-axis quantity on the Axes page of the Graph layout dialog. This has been solved.

- Protocoll window

- Some superfluous empty lines have been removed from the protocoll output.

- The text in the protocoll window was cut-off when the total amount of text to be displayed
      exceeded a certain value. This happened mainly when multiple files were loaded.
      This has been solved.

----------

3 Limitations and known issues

- When RheoWin is installed for a RS100-5Ncm or RS50, the installation program will ask

for the diskette containing the calibration file. If you do not have a calibration file just

click Ok to finish the installation. RheoWin is not able to perform any measurement

using a RS100-5Ncm or RS50 without the correct calibration file !

Please note that you can copy the calibration file into the RheoWin Drivers directory

manually after the installation was completed.

- Most languages (in the language file) apart from German and English are not 100% complete.

An updated language file will be made available on the RheoWin web-site.

----------

4 Contact Thermo Electron

If you have any questions and or suggestion regarding HAAKE RheoWin

please send an e-mail to:

info.ctd.de@thermo.com

----------

This file was last updated on: 07.02.2006 by Jint Nijman