MARS II Firmware C-Box 13.02.001 
			   DSP    11.01.000

Release date: 03 september 2013

Below is a list of things that are new and/or changed in this version compared with 
the previous version. The firmware for the DSP and M-Box have not been modified.

Please note: This version is only compatible with RheoWin 4.41.0000 or newer, that

means this firmware version is NOT compatible with older RheoWin versions!

C-Box

1) Under certain circumstances the angle reset function (at the start of an element) still did 
    not work correctly (see the previous version). This has been solved.

----------

MARS II Firmware C-Box 13.02.000 
			   DSP    11.01.000

Release date: 05 august 2013

Below is a list of things that are new and/or changed in this version compared with 
the previous version. The firmware for the M-Box has not been modified.

Please note: This version is only compatible with RheoWin 4.41.0000 or newer, that

means this firmware version is NOT compatible with older RheoWin versions!

C-Box & DSP firmware

1) The (much) improved CD-Mode of the MARS III has been transfered to the MARS II.

- The CD-Mode has been improved in many ways and now gives better, smoother and
       thus more reproducible results for G’, G”, etc.

- When the max torque of the drive motor is reached in CD-Mode (during curing for example),
      the firmware will automatically switch to CS-Mode using the highest torque value. 
      The firmware will also switch back to CD-Mode when possible.

2) Under certain circumstances the angle reset function (at the start of an element) did 
    not work correctly. This has been solved.

3) The firmware now supports upto 60 steps for the RheoWin MultiStep element.

RheoWin 4.41.0000 or newer is needed for this.

----------

MARS II Firmware C-Box 13.01.000

DSP    10.07.000

M-Box 12.04.000

Release date: 15 november 2012

Below is a list of things that are new and/or changed in this version compared with 
version C-Box 13.00.000, M-Box 12.03.000. The firmware for the DSP has not been modified.

This version is compatible with RheoWin 4.30.0012 or newer, that means this firmware version is NOT compatible with any older version!

RC-Box firmware

1) In all previous firmware versions, when using the AutoTension function activated in a
    RheoWin oscillation element, the normal force control was only active in between two
    oscillation periods. This has been changed. 
    The normal force control loop is now working continuously i.e. independ from the oscillation 
    control and measurment.
    This also means that activating the AutoTension function in the Set-Element or in any 
    Osc-Element will result in exactly the same continous normal force control.

2) A problem that occured while switching from AutoTension to ThermoGap and vice versa 
    has been solved.

M-Box firmware

1) Lift control has been improved:

- Finding the zero position now takes less time.

----------

MARS II Firmware C-Box 13.00.000

DSP    10.07.000

M-Box 12.03.000

Release date: 02 august 2010

Below is a list of things that are new and/or changed in this version compared with 
version C-Box 12.09.000, DSP 10.07.000. The firmware for the M-Box has not been modified.

This version is compatible with RheoWin 4.0 or newer, that means this firmware version is NOT compatible with RheoWin 3.61.000x and older versions!

1)	Compatibility with UTMC

Starting from this firmware version the MARS is compatible with the new UTMC control box which is needed for the new TM-xx-x temperature control modules.

----------

MARS II Firmware C-Box 12.09.000 
                                DSP    10.07.000

M-Box 12.03.000

Release date: 18 august 2009

Below is a list of things that are new and/or changed in this version compared with 
version C-Box 12.06.000, DSP 10.07.000. The firmware for the M-Box has not been modified.

This version is compatible with RheoWin 4.0 or newer, that means this firmware version is NOT compatible with RheoWin 3.61.000x and older versions!

1)	Discontinuous raw sine-wave data in oscillation CS and CD-AS mode (in RheoWin 4)
Due to a modification in the timing/sequence of the communication between RheoWin and the rheometer firmware in RheoWin 4 (relative to RheoWin 3.xx), the raw sine wave data was not transfered correctly. As a result one would see discontinuous sine waves in RheoWin 4 like in the graph below. This was a ‘cosmetic’ problem only and did NOT influence the G’, G”, etc. data at all.
(We never actually saw this problem with the MARS but theoretically it could happen)

2)	Erroneous phase angle values under extreme conditions
When measuring samples that are almost completely elastic (phase angle close to 0.0 °) at (very) small deformations the measured phase angle would vary between the correct value (around 0.0 °) an much larger and wrong values (around 60 °, 90 °, 36°, or 72 °).

----------

MARS II Firmware C-Box 12.06.000

DSP    10.07.000

M-Box 12.03.000

Below is a list of things that are new and/or changed in this version compared with 
version C-Box 12.03.000, DSP 10.03.001 and M-Box 12.02.000.

This version is compatible with RheoWin 3.61.0005 and RheoWin 4.0 or newer, that 
means this firmware version is NOT compatible with RheoWin 3.61.0004 and older versions!

1) A completely new control loop for CR-Mode with the following features:

a)	Very low (down to 10-9 rpm) rotational speed for almost all samples.

b)	Very fast control (t-set < 10 to 20 ms) of the CR mode for rpm values > 10-4 rpm.

c)	The control can be adapted to (almost) any sample by changing some parameters according to some simple rules.

2) A completely new control loop for CD-Mode with the following features:

a)	Stable deformation control of all samples, from purely elastic spring like materials, low viscous gel like materials, to (very) high viscous materials.

b)	automatic switching to CS-Mode when a user defined torque has been exceeded.

c)	automatic switching to CS-Mode when a certain motor temperature has been exceeded.

d)	The control can be optimized for a faster response (when needed) by changing some parameters according to some simple rules. Please note that this feature (changing parameters) is NOT needed for normal operation.

3) Some smaller optimizations for the lift control have been implemented.

4) Initialization and lift-movement

In normal operation mode the lift will move upwards after switching the power on. In order to prevent damage the lift will NOT move when for the last measurement (Job or Monitor) the SER-Tool was used. This security function can only be disabled by running a measurement with another measuring geometry before switching the MARS off (and on again).

Alternatively the lift movement after power up can be suppressed entirely (i.e. independent of the which measuring geometry was last used) by setting DIP switch no. 4 on the back of the MARS control box in the position pointing to the left side of the box (seen from the back of the box).

The drive motor initialisation after power up can surpressed by setting DIP switch no. 3 on the back of the MARS control box in the position pointing to left side of the box (seen from the back of the box). Supressing the drive motor initialisation will disable the absolute angle signal of the rheometer and thus the MSC functionality. It is NOT recommended to us this option.

Per default the four DIP switches should be in the position pointing to the right side of the box (seen from the back of the box).

----------

MARS II Firmware C-Box 12.03.000

DSP    10.03.001

Below is a list of things that are new and/or changed in version compared with 
C-Box version 12.03.000

This version is compatible with RheoWin 3.60 or newer, that means this firmware

version is NOT compatible with RheoWin 3.50 and older!

- The measured torque signal during a MSC measurement was more noisy then it used 
  to be before the last firmware update due to modification elsewhere in the firmware.

This has been modified.

----------

MARS II Firmware C-Box 12.03.000

DSP    10.03.000

Below is a list of things that are new and/or changed in version compared with 
C-Box version 12.02.000

This version is compatible with RheoWin 3.60 or newer, that means this firmware

version is NOT compatible with RheoWin 3.50 and older!

- There was a problem with the communication between the C-Box and the UTC
  control box. This has been solved.
- Some improvements have been made in the calculation of the angular speed in CS-Mode.

- Some smaller bugs in CR-Mode have been fixed.

- New routines for the RheoWin MulitStep Elements have been implemented (see the 
  RheoWin 3.60 fo more information).

----------

MARS II Firmware C-Box 12.02.000

Below is a list of things that are new and/or changed in version compared with 
C-Box version 12.01.000

This version is compatible with RheoWin 3.50.0011 or newer, that means this firmware

version is NOT compatible with RheoWin 3.50 and older!

- The LED at the back of the C-Box that indicates the trigger function did not work 
      under certain circumstances, this has been solved.

- Switching the MARS off (very) shortly after switching it on could result in the instrument 
      not working properly anymore (wrong motorconstants). This has been solved.

- Some new functionality for the RheoScope High Temperature version has been implemented.

- The oscillation frequency at which the motor frequency is changed can now be modified

by an entry in the RheoWin drivers.ini file. This is for very special applications only.

- There were problems with the Axial Ramp element, these have been solved.

----------

MARS II Firmware C-Box 12.01.000,

M-Box 12.02.000

Below is a list of things that are new and/or changed in version compared with 
C-Box version 12.00.000 and M-Box version 12.01.00

This version is compatible with RheoWin 3.40 or newer.

- Some changes in the lift-control with regard to the RheoScope module and the Construction
  Material Cell. In both cases there were problems determining the gap zero position. 
  This has been solved.

----------

MARS II Firmware C-Box 12.00.000,

M-Box 12.01.000

Below is a list of things that are new and/or changed in version compared with 
C-Box version 11.03.020 and M-Box version 12.00.00

This version is compatible with RheoWin 3.40 or newer.

- Moving the lift  to the measurement position can now optionally (this is an option in the 
  RheoWin Lift-Element) be combined with setting  the optimal angular position and an 
  automatic optimization of the MSC correction.
- The oscillation CD-Mode control loop has been improved. This mainly effects the 
  measurement of low viscous but highly elastic samples like gels and (semi-) solid elastic 
  materials. 
- The mean absolute angle (at which a data point was taken) is now available for
  all oscillation data.
- Functionality for controlling the RheoScope Modul has been added.
- Some bugs regarding the use of Thermogap and AutoTension during oscillation 
  measurements have been fixed.
- Some minor bugs have been fixed.

----------

MARS II Firmware C-Box 11.03.020,

M-Box 12.00.00

Below is a list of things that are new and/or changed in version compared with 
C-Box version 11.03.000 and M-Box version 11.01.000.

This version is compatible with RheoWin 3.30 or newer.

- Switching from AutoTension to ThermoGap now works ok.

- The normal force signal is now also properly recorded when ThermoGap is running.

- The normal force control loop is now faster.

- The internal sequence controller now doesn't wait indefinite anymore for the normal force 
  control loop to reach the set normal force. After 15 seconds the next oscillation data point
  will be taken. This waiting time can be changed by modifying a MARS motor constant by 
  using a terminal program (maybe we will make this settable from RheoWin).

- When the CTC is mounted to the MARS but not used (that is: not switched on) but a fluid
  Temperature Control Unit or the UTC is connected to the MARS (and switched on) the lift
  will now automatically stop at middle position (and not move to the top postion from which
  you have to get it down using the two-button-mode on the MARS front panel).

- Controlling a TC501 or TC81 while connected to the MARS C-Box (Remote connection
  in RheoWin) caused problems. These have been solved.

Please note:   It's very important to use version 1.2.0.0 of the Thermo Haake 
                        Firmware Updater (which is part of the RheoWin 3.30 or newer 
                        installation) when updating the MARS firmware!!!!
                        Using an older version of the Firmware Updater will cause problems!

----------

Contact Thermo Fisher Scientific

If you have any questions and or suggestions regarding the MARS (firmware)

please send an e-mail to:

support.mc.de@thermofisher.com

----------

This file was last updated on: 09.09.2013 by Jint Nijman.