RS6000 Firmware Microcontroller  03.03.000 
                                DSP                  10.10.000

Release date: 10 september 2012

Below is a list of things that are new and/or changed in this version compared with 
version Microcontroller 03.03.000 and DSP 10.10.000.

Please note: This version is compatible with RheoWin 4.30.0011 or newer!

DSP firmware

1) Oscillation in CS-Mode with (very) low torques has been improved by implementing the 
    automatic active drift-compensation routine from the MARS III.

RS6000 Firmware Microcontroller  03.03.000 
                                DSP                  10.09.001

Release date: 26 april 2012

Below is a list of things that are new and/or changed in this version compared with 
version Microcontroller 03.02.000 and DSP 10.09.000

Please note: This version is compatible with RheoWin 4.30.0011 or newer!

Microcontroller and DSP firmware

1) Improved zero-point finding using normal force (like MARS), i.e. NOT by using rotation 
    and measuring the torque anymore.

2) Lift control logic very similar to that of the MARS, that means

- a more precautious downward movement

- high(er) normal forces only occur at contact position after switching from a shorter

to a longer geometry.

RS6000 Firmware Microcontroller  03.02.000 
                                DSP                  10.09.001

Release date: 28 september 2011

Below is a list of things that are new and/or changed in this version compared with 
version Microcontroller 03.00.000 and DSP 10.09.000

Please note: This version is compatible with RheoWin 4.30.0011 or newer!

Microcontroller and DSP firmware

1) A problem with the lift control and zero point finding in combination with the D70/300 and D400/300 pressure cells has been solved. For installations which do not include this pressure cell, there is no need to update the system with this new firmware version.

2) The ThermoGap control loop will now also correct the gap for changes due to a changing ambient temperature.

3) The ThermoGap control loop has been improved regarding the resolution and the response time.

4) The RheoWin AutoTension functionality did not work when ThermoGap was active, this has been solved.

5) The Normal force measurement has been improved regarding the response time.

RS6000 Firmware Microcontroller  03.00.000 
                               DSP                   10.09.000

Release date: 23 september 2010

Below is a list of things that are new and/or changed in this version compared with 
version Microcontroller 02.22.001 and DSP 10.07.000

Please note: This version is compatible with RheoWin 4.xx.xxxx or newer!

Microcontroller and DSP firmware

1) Under certain circumstances there was a problem when switching from CS-mode to 
    CR-mode, this has been solved.

----------

Contact Thermo Fisher Scientific

If you have any questions and or suggestions regarding the MARS (firmware)

please send an e-mail to:

support.mc.de@thermofisher.com

----------

This file was last updated on: 26.04.2012 by Jint Nijman.