VTiQ Air	Firmware 	V1	02.02.002
 		V2	02.05.000

Release date: 23 november 2020

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

- Under certain circumstances the viscosity value was not correctly displayed on the instruments 
  touchscreen when the viscosity unit was set to P (=Poise) or cP (centiPoise). This has been solved.
  Note: The visosity value displayed in RheoWin was always correct.

VTiQ Air	Firmware 	V1	02.02.000
 		V2	02.05.000

Release date: 03 may 2018

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

- Under certain circumstances (sample with high G* value) the reported torque value in OSC-mode
  was not correctly calculated. This has been solved.
  Note: All other measurement values were correct.

VTiQ Air	Firmware 	V1	02.02.000
 		V2	02.04.000

Release date: 30 october 2017

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

- The data acquisition for rotational measurements has been improved.

- In manual control mode setting a shear rate or shear stress value selected from a row 
  of values defined with RheoApp did not work correctly. This has been solved.

VTiQ Air	Firmware 	V1	02.01.000
 		V2	02.02.000

Release date: 12 october 2016

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

- The quality of the data from oscillatory measurements has been improved (improved
  routines for phase angle calculation, compliance correction, etc.)

- The CR-control loop has been improved and now offers a (much) quicker response time for 
  high(er) viscous samples.

- The bearing friction correction now also corrects for changes in the motor temperature.

- The data acqusition during CR and CS rotational ramps has been improved.

- On the Login screen the default user was not disabled when other users are defined (using
  the RheoApp software). This has been solved.
  Please note: This firmware version must be used with RheoApp version 2.4.2.9 or higher.
- The cone angle value for cone&plate measuring geometries was not displayed correctly.
  This has been solved.

- The temperature of the Peltier heatsink is now continuously monitored. When the heatsink
  temperature reaches an unexpected value (due to insufficient cooling), the Peltier 
  temperature control will be halted and a message will be displayed on the touchscreen.

- The option to automatically activate the Peltier temperature control after the instrument is 
  switched on can now be activate from the touchscreen user interface in 
  the Configuration > Device settings menu.

- It was not possible to enter a geometry factor value higher than 1e+6 using the touchscreen 
  user interface. This has been solved.

VTiQ Air	Firmware 	V1	02.00.000
 		V2	02.00.000

Release date: 4 may 2016

These are the first release versions for the VTiQ Air.

----------

Contact Thermo Fisher Scientific

If you have any questions and or suggestions regarding the MARS (firmware)

please send an e-mail to:

support.mc.de@thermofisher.com

----------

This file was last updated on: 25.11.2020 by Jint Nijman.