HAAKE RheoWin Version 3.50.0012

Contents of this document

1 Downloading and installing

1.1 Download the RheoWin files

1.2 Installation of RheoWin 3

1.2.1 Installation

1.2.2 RheoWin and Windows Users (Administrators, PowerUsers, Users, etc.)

1.3 RheoWin 3 license key

2 What's new

2.1 Version 3.0

2.2 Version 3.1

2.3 Version 3.11

2.4 Version 3.12

2.5 Version 3.14

2.6 Version 3.20

2.7 Version 3.21

2.8 Version 3.22

2.9 Version 3.22 build 4

2.10 Version 3.23

2.11 Version 3.23.0003

2.12 Version 3.30

2.13 Version 3.40

2.14 Version 3.50

2.15 Version 3.50.0011

2.16 Version 3.50.0012

3 Limitations and known problems

4 Contact Thermo Electron Corporation

============================================================

1 Downloading and installing

This is a fully executable version of RheoWin. Please note that in order to be able to

install RheoWin on your computer you need a RheoWin 3 License Key.

RheoWin 3 currently supports the following HAAKE rheometers and viscometers:

MARS II, RheoStress 600, RheoStress 300, RheoStress 150, RheoStress 80, 
RheoStress 75, RheoStress 100-1Ncm, RheoStress 100-5Ncm, RheoStress 50, 
RT20, VT550, VT6L, VT6R, VT7L, VT7R, VT7L-plus, VT7R-plus, RotoVisco 1, 
RheoStress 1 and RheoScope 1.

RheoWin 3 does not support older instruments like the RT10, VT500, VT501 and the

RV20+RC20/RV30 with CV20, CV100, M5 or M10.

Although RheoWin 3 runs under Windows 98, Windows 98 SE, Windows ME,

Windows NT4.0, Windows 2000 and Windows XP, we strongly recommend to run 
RheoWin 3 on Windows 2000 or Windows XP since Microsoft discontinued the support 
for Windows 98, Windows 98 SE, Windows ME and Windows NT4.0.

RheoWin 3 will NOT run under Windows 3.x, Windows NT3.x and Windows 95 !

IMPORTANT: In order to be able to make RheoWin 3 compliant with FDA 21 CFR part 11 
it is required that RheoWin3 runs under (Windows NT 4.0, not recommended, see above) Windows 2000 or Windows XP.

1.1 Download the RheoWin files

Download the following two *.ZIP files and copy them to a single (temporary) directory

on your hard drive.

Files : Rhw350_1.zip and Rhw350_2.zip

(On www.thermo.com/mc these two files will have different names)

Then unzip each of the *.ZIP files using any suitable unzip program. In case you do not

have an (UN)ZIP program you can download a freeware ZIP program called FreeZip from the

following Web-site : "http://www.ozemail.com.au/~nulifetv/freezip" .

Because of the size of the *.cab files it is not possible to copy the installation files

to floppy disks.

1.2 Installation of RheoWin 3

RheoWin 3 can be installed parallel to an existing RheoWin 2.xx installation if needed. Please

note that RheoWin 3 can read RheoWin 2.xx data, job and page files (backward compatibility),

but that RheoWin 2.xx can NOT read RheoWin 3 files !!

RheoWin 3 should never be installed in the same directory as RheoWin 2.xx !!

The installation program of RheoWin 3.4 offers the user the possibility to transfer all

settings (apart from the Unit settings) from a previous installation of RheoWin 3.xx. Please

refer to paragraph 2.7 and 3 below for details.

In contrast to what is written in paragraph 2.7 in the case of an update from version 3.21
(or 3.21.x) or 3.22 to 3.23 (and only in that case) the installation program of version 
3.23 and higher does NOT have to be started before uninstalling the previous version!!

1.2.1 Installation

Here's how to install RheoWin:

1) Please note that (Windows) Administrator rights are needed to install RheoWin!!!

2) Deinstall (remove) any previous RheoWin 3 version first (if applicable):

Click the Windows Start button, then select Settings, then select Control Panel, 
    in the Control Panel select Add or Remove Programs, in the list of programs select 
    HAAKE RheoWin and then click on the Change/Remove button and follow the instructions.

3) From a directory on your harddisk in which you unzipped the *.ZIP files :

To install RheoWin from harddisk go to the (temporary) directory in which

you unzipped the *.ZIP files after downloading. Please make sure that you

have the following 7 files in that directory :

HAAKE RheoWin.msi, Setup.exe, Data1.cab, 
              1031.mst, 1033.mst, 1034.mst, 1036.mst, 1043.mst

To start the installation please run the Setup.exe program.

4) From the RheoWin CD :
    This CD is "self-starting", just click on "Install HAAKE RheoWin 3" to start the installation.

In case the installation screen does not come up automatically you can start it by

double-clicking on the file "index.htm" in the Windows explorer, this will show the 
    installation screen in your web-browser.

You can also start the installation by run the setup.exe program from the RheoWin

directory on the CD

The installation program will guide you through the installation. Please make sure that you

have a RheoWin 3 License Key (see below) readily available since the installation program

will ask for this.

It is strongly recommended to reboot the PC after installing RheoWin!

1.2.2 RheoWin and Windows Users (Administrators, PowerUsers, Users, etc.)

Windows users with the default Administrator or PowerUser rights will be able to run RheoWin without the need for any modifications.

For Windows users with the default Users rights an Administrator has to make sure that those users have read and write (!) access to the RheoWin 3 main directory and all its subdirectories, as well read and write (!) access to the 
	HKEY_LOCAL_MACHINE\Software\Thermo Electron\RheoWin3 
key in the windows registry.

1.3 RheoWin 3 Licence Key

In order to be able to install this version of RheoWin on your computer you will need a 
RheoWin 3 License Key. A Key-Diskette from a previous version of RheoWin will NOT 
enable you to install RheoWin 3.

The keys for RheoWin 3 consist of a company name, serial number and the key code. You

will have to enter this information manually during the installation. When you receive the key

per e-mail this information will be stored in a *.key or *.txt file which can be opened and viewed by using any text-editor (Notepad, Write, Word, etc.)

In order to be able to install the new tools for FDA 21 CFR part 11 compliance, a separate

second License Key is needed. A separate key is also needed to install the new TTS (Time

Temperature Superposition) tools.

License Keys for RheoWin 3 are available on request only !

You can request a RheoWin 3 license key from the RheoWin download page on

our web-site at www.thermo.com/mc (see under ressources at the bottom of the page).

----------

2 What's new

2.1 Version 3.0

- Installation program

- RheoWin 3 uses completely new License Keys. The Key-Diskettes from RheoWin 2.xx 
    will not work with RheoWin 3. 
  - The new License Keys consists of the customers company name, a serial number and

a 24 digit code, all of which has to be entered in a special dialog during installation.

- The installation program is now based on the Windows Installer (MSI) technology. 
  - To to add other rheometers to an existing installation run the installation program again.
    The installation program will then start in the so-called Maintenance mode, from the 
    Maintenance mode welcome screen select Modify, you will need the appropriate License 
    Key to continue and be able to install additional rheometers.

- To remove RheoWin from a computer run the installation program again.The installation

program will then start in the so-called Maintenance mode, from the Maintenance mode 
    welcome screen select Remove to de-install the software.
    Do NOT try to remove RheoWin from a computer by deleting the RheoWin directories and 
    deleting entries in the Windows registry !

- User Management System
  - The User Management System is completely new and designed to comply with the
    "FDA 21 CFR part" regulations for a closed system.
  - The User Management System can be started as a separate program (UMSDlg.exe in the

RheoWin directory) or from the JobManager and DataManager Configuration menu.

- Using the User Management System an administrator can define users (user accounts)
    and assign privileges to each user. To make assigning priviliges easier, there are four

pre-defined privileges groups. These groups can be modified, it is also possible to create

new groups.
  - There is wide range of possibilities to set the properties of user passwords and user

accounts.

- The user management system can be disabled by activating the default account on the

Administration tab of the user management system dialog.

- JobManager

- All the Element icons are now displayed in three groups (General, Measurement,      
    Evaluation) but in one window (instead of in three separate windows). 
    The sorting order of the Element icons is now always the same.

- The JobEditor has been redesigned :
    - To remove an Element you have to drag the icon on the litter can

- New Element can be inserted by slecting Insert new from the right-mouse button 
      popup window

- Multiple Elements can be selected by pressing the Crtl or Shift key when clicking the 
      left mouse key (standard Windows method for multiple selection from lists).
    - Multiple selected Elements can be copied to another position in the same or another

Job by Drag&Drop

- The Rheometer, Sensor and Thermocontroller selection is now directly integrated in the
      JobEditor main dialog. Clicking the right mouse button will show a small popup menu with

the following commands Edit, Test and Diagnosis. These commands provide a direct link 
      to the corresponding commands in the DeviceManager.

- The JobExecute and JobDisplay windows are now combined into one window (JobRun

window). The display part of this window has three pages, the first displays a standard

RheoWin Page (graph, table and protocoll), the second one a second graph and the third
      one a second table

- The extended version of manual control window now contains four pages, i.e. two

separate graphs and two separate table.

- DataManager

- The page window is now a so called "splitter window", this means that the size of the
    graph, table and protocoll area inside the window can be changed by clicking on the 
    dividing line and draging that line to a new position.

NEW Optional 21 CFR Part 11 tools

The new and optional RheoWin 21 CFR Part 11 Tools supply Audit Trail and Electronic 
Signature functionality for RheoWin 3. Combined with the new User Management System 
this makes RheoWin "technical compliant" with the FDA's 21 CFR Part 11 requirements.

- Audit Trail

- The Audit Trail records all changes made (and saved) in the User Management System, 
    Job files (*.rwj), Data files (*.rwd) and Page files (*.rwp).

- The Audit Trail is stored and maintained in a database file using an SQL Server.
  - By default the SQL server is connected to an Microsoft Access database file

HAAKEAuditTrail.mdb which is stored in the RheoWin directory.

- An entry in the Audit Trail contains the date and time of the change, the name of the 
    operator who performed the change, the name of the record (file) that was changed 
    and a detailed listing of the parameters that were changed as well as a reason for 
    the changes.

- Electronic Signatures

- Up to five Electronic Signatures can be applied to Job (*.rwj), Data (*.rwd) and

Page (*.rwp) files.

- The number of signatures that are needed to "lock" a file can be set on the

Security options page of the Preferences dialog in the Configuration menu.

- A file that is locked can not be modified anymore.

- Signature meanings ("Approval", "Review" etc.) are user defineable.

- Job files can signed the following ways :

- by selecting Sign job from the JobManager file menu

- by clicking on the Key icon in the Icon bar of the JobManager

- by right clicking on the upper part of the Job Editor

- Data files can be signed from the File menu in DataManager using the Sign file command

- Page files can be signed from the Page menu of the File menu using the Sign page 
    command

2.2 Version 3.1

Below is a list of things that are new compared with version 3.0 of RheoWin.

- General

- In previous versions of RheoWin data could not be saved in directories where the

security settings do not allow files to be deleted. This has been solved.

- The quantity "kinematic viscosity" (viscosity/density) has been added.

This quantity is available in the results of all ROT elements.

- The font file that is supplied with RheoWin is now based on a Uni-Code font file.

- AuditTrail

- The software installation is now logged in the AuditTrail database file.

- Rheometer drivers

- This version includes drivers for the new HAAKE VT7R-plus and VT7L-plus.

- DeviceManager

- UTC
    - A second temperatur offset table, for the upper heater of the TC501, has been added.

- Phoenix circulator
    - There were problems reading the actual temperature when the circulator was connected
      directly to the PC. This has been solved.

- Inertia calibration

There was a problem with the inertia calibration for measuring geometries.

This has been solved.

- JobManager

- File select dialog

Under certain circumstances the File Select dialog was not automatically displayed

in the foreground (and hidden by the JobController). This has been solved.

- RheoWizard
    Some problems with the RheoWizard that would cause RheoWin to crash have

been solved.

- JobDisplay window
    The definitions of the (second) graph and (second) table on the second and third page 
    of the JobDisplay window can now be defined and saved in the same way as the graph

and table of the first page of this window.

- Crossover evaluation Element
    The crossover calculation now also works on data measured with a CD-amplitude sweep.

- Regression evaluation Element

The regression calculation is now performed on the quantities in the units as they are

displayed on on the relevant axes in the graph. 
    In previous versions of RheoWin the calculation was always performed on the quantities

expressed in their basic SI units even if the units displayed on the axes where not

in the basic units.

- JobManager / DataManager

- Electronic signatures
    When the current logged in user did not have the privileges to apply signatures

another user (who had that privilige) could not apply a signature either, because the

signature dialog was not accessible.

This has been changed.

Users with the appropriate priviliges can now always apply their signature independent

of the status of the currently logged-in user.

- DataManager

- Graph definition dialog

The axes definition page of this dialog has been slightly modified for better useability. 
    The quantity selection box is higher and wider. The (corresponding) X-axis quantity is

now listed behind the Y-axis quantity, this makes the handling of graphs with two

(or more) quantities on the X-axis more concise.

- Protokoll and Report

The inertia and damping values of the measuring geometries are now shown in

the Protokoll/Info display.
    The Measurereport.lst and Standardreport.lst report templates have been changed

and will now show the inertia and damping values.

- Reports

- The newest version of the report tool, List&Label version 10, has been integrated
    in RheoWin. Compared with the previous version that was used, this version offers a
    range of new possibilities. In RheoWin DataManager reports can now be saved in the
    following formats : PDF, HTML, Excel, RFT, TXT, TIFF, JPG etc.

2.3 Version 3.11

Below is a list of things that are new compared with version 3.1 of RheoWin.

- DeviceManager

- Inertia calibration

There was a problem with the inertia calibration for rheometers.

This has been solved.

- Import, Export and Replace buttons

Under Windows 98 the Import, Export and Replace buttons in the DeviceManager

did not work. This has been solved.

- MSC and MFC calibration

The determined MSC and MFC tables were not properly transfered to the rheometer. 
    This has been solved.

- JobManager

- Save data file Element

On the Segment page of the Save data file Element there is a new option Save segments 
    separately. When this option is enabled the data from each segment is automatically

saved in a seperate file, i.e. one file per segment. The filenames of these individual files

are automatically generated from the filename XYZ.rwd of the complete file using the

format : Filename is "XYZ" + segmentname + "Segment" + Segmentnumber.rwd. 
    Additionally the normal complete file is saved. The option No data file on

the General page of the Save data file Element is valid for the separate segment files also !

- JobDisplay window
    In the second graph, which displays the main measured quantity during stationary flow

curves etc., the axes were not set correctly. This has been solved.

- Message Element
    RheoWin crashed when the slected soundfile was tested. This has been solved.

- DataManager

- File save dialog

In the File menu there is a new command Save segments separately that allows the user

to save the data from each segment in a separate file, i.e. one file per segment.

The filenames of these individual files are automatically generated from the filename

XYZ.rwd of the orginal file using the format : 
             Filename is "XYZ" + segmentname + "Segment" + Segmentnumber.rwd. 
    These filenames are shown in the Select segments dialog that

pops up when the command Save segments separately is selected. In this dialog the user

can also change the filenames of the individual segments and select the segments that

will be saved (in separate files)

2.4 Version 3.12

Below is a list of things that are new and/or changed compared with version 3.11 of RheoWin.

- DeviceManager

- Thermocontroller TC500 and TC501

- There were some problems on the Temperature page of the DeviceManager dialog

for the TC500 and TC501. This has been solved.

- RheoScope 1 (RO1)

- There were some problems on the Additional data page of the DeviceManager dialog

for the RheoScope 1 (RO1). This has been solved.

- Editors for sensors

- There were some problems on the MSC page of the DeviceManager dialog for

sensors. This has been solved.

2.5 Version 3.14

Below is a list of things that are new and/or changed compared with version 3.12 of RheoWin.

- Installation program

- The installation program now displays some information regarding entering the

registration data (company name, serial number, registration code).

- The installation program now displays some information regarding the handling of a 
    password and username for the AuditTrail database file.

- The installation program will now display a message when the current Windows user does 
    not have administrator level priviliges on the computer.

In this case the installation will be aborted.

- In previous versions of RheoWin 3 the installation program modified a Windows registry 
    entry of RheoWin 2.xx. Because of that patches for RheoWin 2.xx would not work. 
    This has been solved.

- On certain computers RheoWin 3.12 would not run after a succesfull installation.  
    This has been solved

- General

- Menu texts
    A few menu texts were displayed in the wrong language. This has been solved.

- RheoWin Log File

When JobManager and DataManager are executed with the command line parameter 
    /debug, both programs will now write a log file (Rheowin.log) for problem tracing purposes.

- Animations in dialogs

To prevent possible problems with the animations that are used in certain dialogs, the 
    animations *.AVI files, which are stored in the RheoWin3\Animation directory, can now 
    be removed (deleted) without causing problems for the program.

- Starting Jobs from Desktop

Starting a job from the desktop directly (using the jobrun.exe program) ignored the 
    RheoWin User Mangement System, which is a problem for 21 CFR Part11 compliant 
    systems. This problems has been solved, the jobrun.exe program now also uses the
    User Mangement System.

- Device drivers

- All Rheometers
    In RheoWin 3.12 there was a problem with the phase angle in oscillation measurements.

This has been solved.

- RheoScope 1

During a measurement there were no images displayed in the software.
    This has been solved.

- DeviceManager

- Diagnosis

When the diagnosis dialog was closed using the [x] button in the title bar, RheoWin 
    crashed. This has been solved.

- Sensor Inertia determination

When the sensor determination was started with RheoWin controlling multiple 
    rheometers, it was not possible to select the correct instrument.

- JobManager

- Manual Control

The manual contol window now only displays the controls for those modes (CS, CR) that 
    the current instrument actually supports.

- Axial Ramp Element

The axial ramp element did work properly. This has been solved.

- Load File Element

Data from files loaded with this element was not displayed in graphs and tables.

This has been solved.

- Regression Element
    The regression element now also calculates Chi² and R when all model parameters

are fixed.

- Reference Element
    The calculated results from this element were not displayed in the graphs. 
    This has been solved.

- Area Element

The sum value of the different area values was not calculated correctly. 
    This has been solved.

- Crossover Element
    The Crossover element now also calculates the time value for the Crossover point 
    (for OSC Time Curves).

2.6 Version 3.20

Below is a list of things that are new and/or changed compared with version 3.14 of RheoWin.

- Installation program

- The drivers for VT500, VT501, RT10 and RotoVisco 20 and 30 are no longer part of the

RheoWin 3 installation package, i.e. these instruments are not supported in RheoWin 3!

- The following measuring geometries for the VT6 and VT7(plus) have been added to

the RheoWin installation package : TL5, TL6, TL7, TR8, TR9, TR10, TR11 as well

the LCP and all six Helipath measuring geometries (with separate factors for the L and

R versions of the ViscoTesters).

- The main RheoWin registry key has been changed to 
    HKEY_LOCALMACHINE\SOFTWARE\Thermo Electron\RheoWin3
    All RheoWin registry entries are now made under this key.

- The installation program will now check for an already installed (previous) version

of RheoWin. In the case a previous installation is found, the installation program will 
    display a message telling the user to de-install that previous version first.

In this case the installation of version 3.20 will be aborted.
    (For problems with this see paragraph 3 below!!)

- General

- An user manual for the Report editor (List&Label) is now part of the installation package. 
    A German and English language version of this manual can be found in the new RheoWin

Documents directory. 
  - The symbol "delta" was missing in the RheoWin font file. This has been solved.

- Some times the € symbol appeared in a graph for no good reason. This has been solved.

- AuditTrail

- RheoWin would not work with an audit trail database file in MS Access format that was

password protected. This has been solved.

- UserManager

- Error messages displayed by the UserManager are now more specific.

- Device drivers

- Manometer driver
    The driver did not pass the correct pressure value to the software. This has been solved.

- RheoStress 1, RotoVisco 1 and RheoScope 1

The drivers for these instruments did not read the correct temperature from a circulator

(connected to a second RS232 port on the PC) in a ROT measuring element when the

circulator (Phoenix or F6/8, N6/8) was set to "external control" . This has been solved

- RheoStress 600, RheoStress 300

- There was a problem when connecting temperature controllers Remote to

the rheometer. This has been solved.

- MARS, RheoStress 600, RheoStress 300, RheoStress 1, RheoScope 1, RotoVisco 1

- The message "Coefficient not set" could pop up for no obvious reason. 
      This has been solved.

- MARS driver

- Switching between "real" CD-Mode and "AutoStrain" CD-Mode is now performed

by the firmware.

- The information that is supplied by RheoWin for the MARS internal web-pages, about

the current and next running element in a Job, has been extended.

- DeviceManager

- MARS properties dialog
    - Instead of entering the URL or network adress of a MARS instrument manually in

the Interface box on the General tab of this dialog, the software will now search for

all MARS instruments connected to the local network (segment) when clicking

the new Scan button.
    - The lowest torque value that can be used for the MARS is now 0,01 microNm.

- MSC determination

- There was a problem with entering a sensor gap value after selecting a sensor in

the MSC dialog. This has been solved.

- For the MARS the MSC determination has been modified. During the MSC determination

the drive motor will now make (at least) one full turn in clock-wise AND one full turn in

counter-clock-wise direction. 
      This allows for the automatic calculation of the MSC offset value and

the best angular position (at which a measurement should be started).

- MicroTorque calibration (MARS only)

- Using the new MicroTorque calibration the standard torque calibration can be corrected

for small non-linearities in the low-torque range.

- MSC and inertia determination

- When the user tries to start the MSC or inertia determination when the manual control or

a job is running, RheoWin will now warn the user that this will not work.

- Phoenix circulator "offset" problem
    - For the case that a Phoenix circulator is controlled using a second COM interface

the problem with the temperature offset has been solved. When using the take from

previous option the offset value will not be added from element to element anymore,

also RheoWin will display the correct temperature. However the temperature on the

Phoenix display will differ from that displayed from that in RheoWin, a solution for that

will be supplied later.
      Please note : The above mentioned problems are not solved yet for the case that a

Phoenix circulator is connected to the RS600 directly !!! It is therefore strongly

recomended not to use a Phoenix circulator directly connected to a RS600.

- UTC properties dialog

- On the Alarm page of the Temperature page of the UTC properties dialog the user can

now set a upper and lower alarm temperature value at which the UTC will switch off its

heating/cooling unit. It is recommended to set the alarm temperatures 5 degrees

higher/lower then the maximum temperatures used in daily work.

Please note : A firmware update is needed for both the Rheometer (RS1, RS600) and

the UTC for this to work.

UTC      firmware version: 02.03.000
                                        file name: UTC_02_03_000.H86.
                            RSxxx  firmware versions: V1:02.09.053 and V2:02.00.013, 
                                        file name: Rxx_B_02_09_053_dsp_B_02_00_013.h86

- Sensors

- A and M factor for torsional oscillation of solid samples now inlcuded.

- JobManager

- Jobeditor

- When clicking on the yellow "i" in the upper left corner of the JobEditor dialog a window

will appear that displays the min-max parameter values for the current instrument setup.

- The temperature for all elements can now be set/changed at once in one dialog. This

dialog is accessed by clicking on the button with a fuchsia coloured T. In the case the

Job contains one or more temperature ramps of some kind, the dialog offers the

possibility to shift the end temperature according the the change in the start temperature.

- The right mouse key pop-up menu of the Rheometer selection field now contains

a new item Browser when the currently selected rheometer is a MARS.
      By selecting the Browser command RheoWin will open the MARS web-site which is

stored and running inside the MARS instrument itself.

- Monitor/Manual Control

- In the display window with the large fonts, negative values for the normal force where not

displayed. This has been solved.

- When the Monitor (Manual Control) dialog cannot establish a connection to the

instrument this will now be indicated by a blinking bitmap in the upper left corner

of the Monitor dialog.

- By clicking on the new [ >>> ] button right from the text Lift Control in the upper left

corner of the Monitor window, the right hand side part of the Monitor window will now

display the current gap, temperature and normal force in large characters for

easy viewing from a distance.

- The Reset Nf (normal force) function did not work. This has been solved.

- There were some problems with opening and closing the manual control window.
      This has been solved.

- File Load Element

- By including multiple copies (up to 5) of this element in a job, multiple (up to 5) data files

can be displayed in the graph display of the job.

- CS/CR Rotation Step Element

- The Acquisition page of the editor of this element has been modified. 
      1) In the box Step mode the user can now choose from two different modes: The mode 
      Wait for equilibrium at each step provides the same functionality as in previous RheoWin

versions. In the new mode Fixed duration for each step, each step will take exactly as

long as defined by the Step duration, the actual data acquisition (see below) takes places

within this time.

2) In the box Data acquisition mode the user can now choose from two different modes:
      The mode Take average of provides the same functionality as in previous RheoWin

versions. In the new mode Integration time RheoWin will take the average of all the (raw)

data points measured during that time.

The Step mode options are independent of the Data acquisition mode options

and vice versa !

- Information on under which circumstances a measurement value was recorded is now

stored in the new "quantity" D-Info (from Data information). It is stored whether the

equilibrium value or the maximum waiting time was reached, whether the user pressed

the Accept Button etc. etc.

- A bug in the data calculation that could lead to crashes has been removed.
    - Some missing text in certain languages have been added.

- Lift Element

- On the Options page of the Axial page of the Lift Element editor, the option Start

measurement without correct gap now also applies for parallel plate (PP) geometries.
      This option is available for the RheoStress 600, RheoScope 1 and RS150 and usefull for 
      achieving shorter waiting times by not waiting for the exact gap to be reached. 
      (When the correct gap is not reached the software will use the actual gap

for the calculation of the shear-reate etc.)

- The option to have the sensor rotating when the gap is opened is now available for 
      the RotoVisco 1, RheoScope 1 and RheoStress 1. This was already available for the 
      RheoStress 600.

- Check set temperature dialog

The check set temperature dialog that pops up when a Job is cancelled is now

automatically put in the foreground (i.e. in front of all other windows on the screen).

- All Measurement Element
    - The maximum number of datapoints pro element that can be acquired and stored has

been increased from 65.535 to 1.000.000.000. This is mainly for the MARS.

- Rotation Time Element and Rotation Ramp Element

- The timing for the logaritmic equidistant data acquisition has been optimized. Data is

now also acquired for times below 1 second.

- New Execute Element

-  With the new Execute Element any windows program can be started automatically
       as part of a Job sequence.

- Interpolation Element

- Some edit fields in the element editor have been enlarged.

- The result table was not displayed correctly after opening the editor.

This has been solved.
  - Yield-Stress Element

- Under certain circumstances the curvefit lines of the manual mode were not plotted
      correctly. This has been solved.

- DataManager

- Automatic data update
    Using the new "Automatic data update" button on the toolbar DataManager will now

automatically update (i.e. reload) the datafile that was first loaded in the active page

window. This is usefull for monitoring a (long) job running on an other (remote) PC.
    This works as follows: JobManager by default saves the data from a running job in the

autosave.rwd file (in the RheoWin Autosave directory) every 2 minutes and, if setup

accordingly, in a user defined file (by using the File Save Element). By making one of

these two files available for an other PC (by saving it on a network drive for example), a

copy of DataManager running on an other PC can automatically load, show and update

that file and thereby monitor the job process.

- Time Temperature Superposition

Time Temperature Superposition calculation has been added to the Analysis functionality. 
    An extra installation key is needed to be able to install and use this feature.

- Table view

Textinformation about the new "quantity" D-Info (see above) is available by right

clicking on a table entry and selecting Info Q-Channel.

- Loading pages
    - Pages from RheoWin 3.14 (and earlier versions) were not correctly loaded into

RheoWin 3.15.xx. This has been solved.

- Save all
    - When an error occured during the saving of multiple files RheoWin would give the error 
      messages "VisChild save all". This has been solved, RheoWin will now show a list of 
      the saved and not saved files together with a reason why a file was not saved.
  - Saving a graph

- When saving graphs in the WMF format multiple times, the same graph was saved

every time. This has been solved.

- Graph / Table display

- Graph and Table layout dialog

- The Filter settings in both the Graph and Table layout dialog could not be operated using

a mouse. This has been solved.

- There was a problem with selecting segments in both Graph and Table Layout dialog

when scrolling through the segment list. This has been solved

- Dialog handling

- It is no longer possible to open multiple instances of the graph and table

properties dialogs.

- When trying to open a properties dialog that is already open, the dialog will be put

in the foreground instead.

- Opening a right-mouse-key pop-up menu (for modifying the axis scaling) while

the graph properties dialog is open is inhibited, the properties dialog will then be put

in the foreground instead.

- At the end of a Job all graph and table properties dialogs are closed automatically

- Table display

- The code for the table display has been optimized. Table display is much quicker now.

2.7 Version 3.21

Unfortunately a few bugs were discovered in RheoWin 3.20 shortly after it was released.
Version 3.21 fixes these bugs but does not contain any real new functionality.

- Installation program

- At the end of the installation the installation programm know recommendeds to reboot the

PC after the installation is finished. This is mainly to avoid problems with the special font

that is used in RheoWin.

- The installation program now offers the user the possibility to transfer all settings

(apart from the Unit settings) from a previous installation of RheoWin 3.xx to

the new installation. 
    Please note: The settings from a RheoWin 2.xx installation can not be transfered

to a RheoWin 3.xx installation!
  
    In order for this to work, the installation program of version 3.21 has to be started before

the previous version is uninstalled!! The installation program of version 3.21 will now save

settings of the previous version and show a message that the previous version should

be deinstalled before version 3.21 can be installed. After showing that message the

installation program will be aborted. (With the next (future) new version the above will

not be necessary anymore.)
    The user then has to uninstall the previous version. When that is finished, the

installation program (of version 3.21) should be started again. The installation program will

now ask whether the settings of the previous installation should be used for the new

installation or not.

The following settings are transfered when this option is used:

- The RheoWin installation directory! 
       - All settings regarding rheometers, measuring geometries and thermo controllers, i.e. all

settings that can be edited or created within the DeviceManager. This includes things

like temperature offsets, MSC tables, geometry factors, etc.

- All settings made in the Configuration -> Preferences Dialog. These settings are

saved for each user separately!
       - All settings made in the UserManager, i.e. user names, user privileges, etc.

- The chosen user interface language (Saved for each user seperately).

Please note: Currently the Unit setting are NOT saved!!

- Device drivers

- Manometer and Hygrometer driver
    These drivers were not correctly installed in previous versions of RheoWin.

This has been solved.

- DeviceManager

- MSC calibration for RheoStress 600 and RheoStress 300 with Normal Force option
    - The determined MSC (and MFC) calibration values were not transfered to the instrument.

This has been solved.

- MARS, communication log file

- It is strongly recommended not to use the communication log file during normal

measurements. The high data acquisition rate of the MARS results in large amounts

of data being transfered between the MARS instrument and the RheoWin software.

When the communication log file is switched on the data acquisition might get disturbed.

When a measurement is started with the communication log file switched on, the user

will receive a warning and the possibility to switch the log file off for that measurement.

- JobManager

- Save Element
    - When RheoWin is started with the /debug command line parameter, the Save Element
      will now write additional information in the rheowin.log file that might be usefull for

debugging purposes.

- When a job was stopped using the Stop button in the JobController the Save as ... dialog

would react very slow, or not at all.

This has been solved in RheoWin 3.21 build 8 dated 03.05.2005.

- Yield-Stress Element

- In RheoWin 3.20 the curvefit lines of the manual mode were not plotted correctly when

the unit [%] (instead of [-]) was used for the strain. This has been solved.

2.8 Version 3.22

Below is a list of things that are new and/or changed compared with version 3.21 of RheoWin.

- Installation program

- New graph and table templates have been added to the installation.

- Device drivers

- MARS driver

- Error messages from the CD-Mode are now always displayed correctly.

- Circulator and temperature controller drivers

- When controlling any circulator or temperature controller (apart from the UTC) by

RheoWin together with the HAAKE MARS, a timing problem caused the measured
      time values to be wrong (the time seemed to run in slow motion).
      This has been solved.

- MARS driver
    - Under certain circumstances oscillation data could be "damaged" during a Job.
      This has been solved.

- DC30 driver

- When using a DC30 circulator to control the temperature during a temperature ramp
      the temperature did not change as defined. This has been solved.

- DeviceManager

- MARS MicroTorque table
    - The MicroTorque correction is a "fine-calibration" of the drive motor for the torque

range below 10 microNewtonmeter. The correction values are determined using the 
      MicroTorque calibration routine in the DeviceManger, the values are stored in a table 
      in the firmware.
      At the end of each calibration measurement the determined values are also automatically

saved in the file xxxxx.mtc (see below) in the RheoWin drivers directory.
    - The MicroTorque correction table can be downloaded from and uploaded to the
      instrument firmware using the Save and the Load MTC button in the MARS 
      Diagnosis dialog.
      The MicroTorque correction table is always saved in and read from the file xxxxx.mtc.
    - In the filename xxxxx stands for the name of the instrument as displayed in device-

manager (i.e. xxxxx=MARS2 by default).

- MARS TCP/IP connection window

- In previous versions RheoWin used a "hidden" window (visible on the Windows TaskBar

only) to establish a TCP/IP connection with a MARS. The TCP/IP connection now works

without this "hidden" windows.

- MARS network scan

- RheoWin now searches all (sub-) networks to which the PC (on which RheoWin is

installed) is connected for available MARS rheometers.

- Solid clamping tool

- The solid clamping tool can now be defined and selected (in JobManager) as a 
      measuring geometry (sensor).
      Some problems with the actual use of this measuring geometry in a Job have 
      been solved.

- JobManager

- Program stability

- A number of internal routines have been modfied to increase the stability of JobManager.

- Jobeditor / Temperature Set Element

- When setting/changing the temperature for all elements using the button with the fuchsia
      coloured T, the Deviation and Waiting time of the Temperature Set Element 
      were deleted. This has been solved.

- Manual control
    - Below the big character display of the gap, normal-force and temperature the user can

now set a certain normal force for the instrument to apply to the sample.

- Set-Element

- On the Axial page of the element editor dialog the user can now set a certain normal

force for the instrument to apply to the sample.

- Run-Element

- There was a problem with the element editor. This has been solved.

- Saving data files

- In all previous versions of RheoWin data could only be saved once during a Job. 
      This has been changed in order to allow the user to save the data under different 
      names, in multiple directories or in multiple formats.

- Axial-Ramp-Element

- There was a problem when using this element with a RheoStress 1, RotoVisco 1,

RheoScope 1, RheoStress 300 or RheoStress 600. This has been solved.

- Oscillation Measurement Elements

-  There was a problem when running oscillation measurements with a RS150, RS80,
       RS75 or RT20 rheometer when no temperature signal was available.
       This has been solved.

- Oscillation Amplitude Sweep Element

- RheoWin would crash when running this element when 0.0 was entered for the start or
      end value of a strain-sweep and when the distribution was set to Log. 
      This has been solved

- End-Element

- Some extra debug information is now written into the rheowin.log file when RheoWin

is started with the "/debug" command line parameter. This in order to be able locate

a problem with saving data files.

- Reference-Element
    - This Element would "forget" the settings for the absolute bandwith when it was saved

as part of a Job. This has been solved.

- Report-Element

- The printing of a Report from a Job (by using the Report-Element) is now performed by
      calling a separate program (PrintReport.exe) from the Report-Element. The advantage
      of this is that problems with printing a Report can not influence the Job anymore.
      For the user there a no changes in handling etc.

- DataManager

- Recalculate all

- Under certain circumstances DataManager would crash after displaying the message

"wrong index ..." when the Recalculate all button was used after running an analysis

routine. This has been solved.

- TTS Module

- It was not possible to perform a curve-fit (or any other analysis routine) to the shifted

data displayed in the automatically loaded new page. This has been solved.

- Graph / Table display

- Graph layout dialog

- When plotting multiple sub-quantities (all derived from the same quantity), i.e. for

example T, T-motor and T-PT100, on one axis, with the new option Short title on

the Scaling page activated, RheoWin will only show the main quantity (i.e. T in

this example) as the axis title.

- In order to save space for the graph itself the axis titles for the two left y-axis (Y1 and Y3)

and the axis titles for the two right y-axis (Y2 and Y4) can be plotted on one line (instead

of on two lines) when using the new options Connect left and Connect right respectively

on the Scaling page.

- Graph display

- Sometimes graphs were displayed mirrored when RheoWin was running under

Windows 2000 or Windows XP. This has been solved.

- Reports
  - Report templates

- In the report templates GraphAndTable and J-GraphAndTable some text fields were not

correctly defined which resulted in truncated texts.
      This has been solved.

2.9 Version 3.22 build 4

Compared with version 3.22 a few files have been modified to fix some problems.

Pleas note : In case you are installing RheoWin 3.22 build 4 (see the installation 
                    welcome dialog), there is NO need to download and install the files as

mentioned below since these are already included in version 3.22 build 4.

If you already installed version 3.22 you can download these files separately (zipped together into one file NewFiles.zip). After downloading the file, unzip it and replace the original file PrintReport.exe (file size 78 kb) in the main RheoWin 3 directory with the new file (file size 143 kb), also copy the 5 ELM*.DLL files into the RheoWin3\Elmdll directory (this will overwrite the existing files).

- Installation program

- The entries TC501 <---> MARS II and TC81 <---> MARS are now automatically added to 
    the list of thermocontrollers in RheoWin DeviceManager.

- A RheoWin 3 group is now added to the Windows Start Menu for all Windows Users.

- RheoWin 3 desktop icons are now installed for all Windows Users.

- JobManager

- OSC Amplitude Sweep Element
    This element crashed during a job run when 0.0 was entered for the start or end

value of the stress or strain.This has been solved.

- Report-Element

The new report element did not show the evaluation results in the preview window

and on printouts. This has been solved.

- Cross-over Evaluation Element

This element now also calculates the values of the stress, strain and temperature at

the cross-over point where G' = G".

2.10 Version 3.23

Below is a list of things that are new and/or changed compared with 
version 3.22 build 4 of RheoWin.

- Installation program

- Graph and table templates for imported CaBER data have been added to the installation, 
    the user can choose whether these files will be installed or not.

- Graph and table templates for imported MSC and MFC data have been added to 
    the installation.
  - Three new report templates for data export in Excel format have been added 
    to the installation.

- General

- The size of a few dialogs (Rheometer properties in DeviceManager, Graph properties
    dialog, etc.) has been changed to solve text display problems when running RheoWin on
     a PC using a Japanese Windows version.

- Device drivers

- MARS driver
    - Due to some modifications in the MARS driver the RheoWin user interface will now react 
      quicker when the software is communicating with the rheometer (i.e. in manual mode 
      and when a job is running).

- There was a problem in the CS/CR-Rotation Step Element when the MARS driver got an
      invalid motor temperature value from the MARS instrument. 
      This has been solved in the driver.
    - New MARS firmware
      New firmware (version 11.00.00 for both the C-Box and the M-Box) is available for 
      the MARS. This version of RheoWin will work with older MARS firmware, but it is 
      strongly recommended to use the new firmware!!

- UTC driver

- The driver will now automatically take note of which temperature control unit is attached 
      to the UTC box (Peltier or TC501 electrical heater) and use this information in the
      DeviceManager.

- DeviceManager

- The "Extra" temperature channel
    - The temperature stored in the T-extra channel (upper temperature of the TC501, heat
      exchanger temperature of Peltier unit, etc.) sometimes showed 0.0 for no good 
      reason. This has been solved.
  - New MARS firmware update utility 
    - A new firmware update (flash) utility is now part of RheoWin. This utility can be started 
      from the Firmware page in the Properties of MARS dialog (in DeviceManager).

      Please note : This utility only works with MARS firmware version 11.xx.xx and higher !!! 
      
      In order to update the firmware from version 10.xx.xx to 11.xx.xx a special Update Tool
      program must first be uploaded to both the C-Box and M-Box using the WinBSL 3.0 
      program. This is described in detail in the information that comes with the new MARS 
      firmware.   
      The new firmware update utility uses the MARS TCP/IP interface instead of the RS232
      (service) interface and allows a much faster updating (flashing) of the firmware. 
      More information is available when clicking the Help button on the main screen of 
      the new Thermo Haake Firmware Updater utility.

- JobManager

- Job Display Settings
    - Under certain circumstances RheoWin would crash when the setting of the additional
      graph that is available during a Job run were changed. This has been solved.

- Identification Element
    - Changes made to the identification information in the "middle" of a Job are now stored.
      For technical reason the indentification information displayed in the protocoll "window"
      during a Job is NOT updated (the modified information is stored and will be displayed
      correctly in DataManager when the data is loaded)

- Lift Element
    - The Go to measurement position with Autotension routine has been slightly modified 
      (mainly for the MARS2, but this applies to all instrument for which this option is valid).
      The routine will now start when the actual measured Normal force is within 5% of the 
      set value.
    - For MARS only: The normal force will no longer be set back to zero when using the 
      Go to measurement position with the Autotension option. This change is really in the
      MARS firmware and not in RheoWin, so the newest MARS firmware is needed for this!
  - Set Element
    - On the new Thermocontroller page the new option Switch of control enables RheoWin 
      to switch the control loop of a thermocontroller (circulator). This is usefull at the end 
      of a Job that runs unattended.

- Creep and Recovery Elements
    - When using logarithmical data point distribution for the data acquisition, there was a
      problem for times below 1 second. This has been solved.

- Curve-Discussion Evaluation Element
     - The readability of the displayed results has been improved.

- Cross-Over Evaluation Element

- This element now also calculates the Polydispersity index PI according to the equation 
       PI = 10^5 / G-crossover.
    - In this element the x-axis range in which the Cross-Over point is calculated can now be 
      defined by the user (like in most other evaluation elements). This is helpful if there 
      is more than one crossover point since the routine only finds the first crossover point
      within the complete range or the user defined range.

- New Structure Recovery Evaluation Element
    - This new element automatically calculates the "structure recovery" of a sample as 
      measured using a sequence of three measurement elements;
         - Element 1 : Determines the initial structure of the sample 
         - Element 2 : "Destroys" the structure of the sample
         - Element 3 : Measures the structure recovery as a function of time

A measure for the structure can be the viscosity or G' or G", etc.
      Currently the following sequences of measuring elements is allowed (for the Evaluation
      Element to work):

Sequence 1                                 Sequence 2
         - Element 1 :  ROT time curve (CR or CS Mode)     OSC time curve (CD or CS Mode)
         - Element 2 :  ROT time curve (CR or CS Mode)     ROT time curve (CR or CS Mode)
         - Element 3 :  ROT time curve (CR or CS Mode)     OSC time curve (CD or CS Mode)

       The new evaluation element calculates the following parameters:

- The absolute structure recovery at the end of element 3

- The amount of relative structure recovery after a certain time (in element 3)

- The time needed for the structure to recover to certain relative value (in element 3)

- The time needed until G' = G" (in element 3)

- Data Export Element -> Excel + PDF export

- The Data Export Element can now automatically export data in Excel format (using report 
      templates) and export reports in PDF format, this in addition to exporting data in plain 
      ASCII format).
      For exporting data in Excel file format the use of one of the three special report

templates (or modified versions of those templates) is recommended (see below)

- DataManager

- TTS Module

- The RheoWin *.rwd file that contains the shifted data now also contains tan(d).
  - Add text from clipboard
    - When adding a text from the windows clipboard which contained more then 1024 
      characters RheoWin would crash. This has been solved. RheoWin will now display a 
      message that the text is too long and the text will NOT be added to the graph.

- New CaBER data file import
    - RheoWin DataManager can now directly read and import *.cbr files created by the 
      CaBER software. Please note that there are no CaBER analysis routines in RheoWin 
      yet. This functionality is currently for testing only. 
      Please note : CaBER *.cbr files are imported using the normal File -> Open dialog, 
                            you have to select All files (*.*) in the Files of type edit box in order 
                            to be able to see the *.cbr files in the dialog. 
  - New MSC and MFC data file import
    - RheoWin DataManager can now directly read and import *.msc and *.mfc files. A *.msc
      file, which contains the MicroStress correction table, is allways created when running 
      the MSC Calibration routine. A *.mfc file, which contains the MicroForce correction table
      for the Normal Force measurement, is only created when running the MSC Calibration
      routine with a RheoStress 600. The automatically created *.msc and *.mfc files always 
      have the name of the rheometer that was used during the calibration, i.e. mars2.msc,
      rs600.mfc, etc.
      Please note : *.msc and *.mfc file are imported using the normal File -> Open dialog, 
                            you have to select All files (*.*) in the Files of type edit box in order to
                            be able to see the *.msc and *.mfc files in the dialog.

- Graph / Table display

- Graph display dialog

- This dialog now contains a (tab) page called "Markers" when the active RheoWin 
      page contains one or more graphical markers (see the description of the new 
      Structure Recovery Element above). Using the controls on the Marker (tab) page 
      the properties of the graphical markers can be modified.

- Reports
  - Special report templates for Excel export

- RheoWin now comes with three report templates that are optimized for Excel export. 
      Using these templates when exporting to Excel file format, the Excel file will contain
      the first data column in column A, the second data column in column B, etc. 
      This is not the case when using any of the other templates, in that case a lot of 
      empty colums will be inserted to match the "design" of the template as good 
      as possible, which is normally what one wants for excel export.

- Empty lines in Analysis results
    - Some analysis routines (namely the interpolation routine) produced unnecessary 
      empty lines in the results displayed in a report. This has been optimized.

- New Raw Data Viewer

- A program (RWRViewer.exe) for viewing oscillation raw data files (*.rwr files) is now part 
    of RheoWin. This program offers a comfortable way to view the oscillation raw data in

three plots; a plot showing the torque and angle signals (sinewaves) as a function of time,
    a Lissajous plot of the torque signal as a function of the angle signal, and a plot showing 
    the torque amplitude as a function of the angle amplitude. By clicking on the Prev 
    Datapoint and Next Datapoint buttons the user can view the sinewave signals of all the
    datapoints.
    The program can be started by double clicking on its file name in the Windows Explorer 
    or by double clicking on a *.rwr file. When using the File -> Open last measurement 
    command the program will load the *.rwr that was created during the last measurement.
    An example *.rwr file, together with the corresponding *.rwd file , is installed in the
    RheoWin data directory.
    Please note : In order to have RheoWin write *.rwr files (for most oscillation experiments
    performed with a certain rheometer) the following options must be activated:
    - The option Oscillation raw data on the Options page of the rheometer Properties editor 
      in RheoWin DeviceManager.
    - The option Save raw data on the Acquisition page of the oscillation elements (not all 
      oscillation elements have this option).

2.11 Version 3.23.0003

Below is a list of things that are new and/or changed compared with 
version 3.23 of RheoWin.

- General

- We have slightly changed the version numbering. The build number zz, which indicates
    that just one or a few files were changed compared with the previous version, is now 
    included in the version number as the last two digits. 
    So the current version number "3.23.0003" is build number zz=03 of version 3.23.00.

The RheoWin version number now has the format x.yy.ttzz. Here x is the main version
    number, yy the sub version number and zz the build number. For official release versions
    tt is always 00, a test version will have a number tt higher then 00. The build number can 
    be used with both official and test versions.

A version number of this form is needed to comply with the Windows Installer that we 
    are using.

- Installation program

- The Graph and Table templates for imported CaBER data where copied to the main 
    RheoWin3 directory instead of the RheoWin3/Templates directory. 
    This has been corrected.

- Starting a Job from the Desktop
  - When starting a Job, that contains the Report-Element, from the Windows Desktop using
    the Jobrun.exe program, the Job could only be run once. At the second start a message
    would popup that the user doesn't have the rights to start the job.
    This has been solved (by means of a new version of the PrintReport.exe file).

- Memory "leak" in the table display
  - RheoWin version 3.22 and 3.23 suffered from a memory leak in the table display which
    caused RheoWin to crash during very long measurements (depending on the size of
    the PC memory). This has been solved.

2.12 Version 3.30

- Installation program

- With the proper registration key (order number 098-1790) the installation program will 
    now only install the DataManager (and the UserManager).
  - The "solid clamp tool" measuring geometry is now automatically installed as part of an
    installation for the MARS rheometer.

- General

- Version numbering

- The complete RheoWin version number x.yy.ttzz, i.e. including the build number zz,

is now stored in all RheoWin data and job files that are created using this version.

- RheoWin web-site
     - Clicking on Help (in the RheoWin JobManager and DataManager) main menu and

then selecting Internet and then RheoWin Update will now load the new web-site 
       www.rheowin.com in your default internet browser.
  - "Job Information" = "Protocol"

- We used the word "Protocol" for the list of information that is (can be) displayed in both
      JobManager and DataManager on the right side of the graph and table.

In english the word "Protocol" is not appropriate for this, starting from this version we 
      will use the words "Job information" instead.

- Device drivers

- RS150 driver

- Resetting the normal force to zero (in the Set-Element and in Manual Control) now also 
      works for the RS150 on both the internal and external normal force signal (depending 
      on which one is activated).

- MARS driver
    - The MSC calibration is now performed in CR-Mode in air !!!
     That means that there should be no sample (i.e. no Newtonian
     oil) in the measuring geometry during the MSC calibration!

The user can switch between the new CR-Mode method and the old CS-Mode
      method by modifying the drivers.ini file in the RheoWin3 drivers directory as follows:

- For MSC in CR-Mode the [MARS2] part of the drivers.ini should contain the lines:
               [MARS2]

MSCMode=CR

;MSCMode=CS
      - For MSC in CS-Mode the [MARS2] part of the drivers.ini should contain the lines:
               [MARS2]

;MSCMode=CR

MSCMode=CS
      It is strongly recommended NOT to the change the default setting and use 
      the new MSC in CR mode in air only.

Please note: For the MSC calibration to work in CR-Mode, C-Box firmware version
                           UPD_MARS_CBOX_11_02_003_DSP_10_01_008.h86
                           or newer is needed!!

- MARS, RheoStress 600, RheoStress 300, RheoStress 1, RheoScope 1 and RS150

- The drivers for these instruments now check the deviation of the actual gap from

the set gap. See the information under Lift-Element below.

- DeviceManager
  - MTC Calibration for MARS

- The MTC calibration method has been modified and is now fully functional. All new 
      MARS will be delivered with an individual calibrated MTC file for better performance 
      at very low torque values.
  - MSC+MFC calibration for RS600

- It was not possible to finish the calibration routine without transfering the determined
      calibration tables to the instrument, the software didn't allow this. 
      This has been changed
  - Thermal expansion coefficient

- The edit field for the thermal expansion coefficient has been removed from the editors 
      of all rheometers!!

- Instead the thermal expansion coefficient of the system consisting of rheometer 
      plus measuring geometry can now be defined for each measuring geometry individually

in the editor of each measuring geometry. 
      The default value for all measuring geometries is 1.1 mm/°C, this is identical to 
      the old default value for all rheometers.
  - Compliance coefficient for for measuring geometries
    - An edit field for setting the compliance of the system consisting of rheometer plus 
      measuring geometry is now available in the editor of each measuring geometry.

Please note: Currently the compliance correction is only functional for the MARS 
                           and only when using MARS C-Box firmware version
                           UPD_MARS_CBOX_11_02_010_DSP_10_01_008.h86
                           or newer. 
  - Device editors

- After opening any device editor from the DeviceManager the Windows TaskManager
      showed 100% CPU usage (by RheoWin). This has been changed.

- JobRun program

- New command-line parameters

- The Jobrun.exe program (which is used for starting pre-defined Jobs directly, i.e. 
      instead of using JobManager) can now be used with two new command-line 
      parameters. 
      With the command-line parameter /SUBSTANCENAME:text the sample name, 
      which will be stored in the RheoWin data file is set to "text", just as if it were entered
      in the Identification dialogue.

With the command-line parameter /SUBSTANCENUMBER:text the sample number, 
      which will be stored in the RheoWin data file is set to "text",  just as if it were 
      entered in the Identification dialogue.
      The following is a example of how to use this:
           jobrun "/SUBSTANCENAME:sample name"
      Please note the use of "" around the command-line parameter to allow for a sample 
      name that contains spaces.

This feature will be used to start RheoWin Job's directly from within a LIMS systems.

- JobManager

- Job window
    - In the Job window the Job list was not always correctly sorted alphabetically (depending
     on the operating system of the PC or network server). This has been solved.
  - Manual control / Monitor Mode

- The default settings for applying a shear rate (angular speed) or shear stress (torque)
      below the graph in the Monitor window has been changed. For all instruments that 
      can be run in CS-Mode, the default mode is now CS-Mode and not CR-Mode anymore.
      This because setting an angular speed of 0.0 will actually activate the speed control 
      loop (which is not expected by most users), while applying a torque of 0.0 will virtually
      not activate any control loop and thus keep the drive motor "free".

- It was possible to open the manual control dialogue and control the instrument when a
      job was already running. Controlling the instrument from the manual control dialogue is 
      now not possible anymore, however the monitor graph/table will still display the values

measured by a job (after clicking the Start button).

- End of Job, log file

- At the end of a Job the communication log file was deleted when it had reached its
      maximum size. This is not the case anymore.

- Element editors

- After opening any element editor from the JobEditor the Windows TaskManager
      showed 100% CPU usage (by RheoWin). This has been changed.

- Save Element

- Some additional debug-output is now written in the rheowin.log file when JobManager 
      is started with the /debug commanline parameter.

- Lift Element

- On the Options page of the Axial page of the Lift-Element editor there is a new option
      "Mark data point when gap deviation >". When this option is activated all data points
      for which the actual gap deviates more then a certain value from the set gap will 
      marked in the Q-channel.
      The Q-channel is bit-coded. If the bit with value 64 is 1, the deviation exceded the 
      tolerance value. By using the context menu (right mouse key pop-up menu) in the
      RheoWin table the user can check the value of the Q-channel.
  - Set Element

- On the Parameter page the Set angle position functionality has been slightly changed.
      The waiting time now also can and should be used when putting the drive motor axis
      in the optimal position. The drive motor will now rotate to that position slowly and wait
      for a few seconds to accurately arrive at that position.
    - On the Axial page the text Set Tension has been changed to the more appropriate 
      text Normal force control. Using the new Off option the normal force can now also be 
      switched off.
    - On the Thermocontroller page the option Switch of control will now also switch off the
      pump of a circulator when it is controlled by RheoWin. 
      When this option is used as the last element in a job that "controls" the thermo-
      controller (circulator), the Job will not check and ask about setting a temperature at 
      the end of the job (other wise the control loop and pump would be switched on again).

- Temperature Set Element

- When running two jobs for two different instruments at the same time it was not (easily)
      possible to see which temperature set dialog "belonged" to which instrument. This has 
      been solved, the name of the instrument is now shown in the title bar of the dialog.

- CS/CR Rotation Step Element

- The use of the mode Fixed duration for each step (in the box Step mode) in combination
      with the mode Integration time (from the box Data acquisition mode) would result in 
      slightly wrong measuring values.
      This has been solved.

- OSC Time, OSC Temperature Ramp and OSC Temperature Step Element
    - In these element a limiting maximum torque for the CD-Mode (AutoStrain) can now 
      be defined.

- CS/CR Rotation Ramp Element

- A new break criteria Break at maximum in Fn has been implemented in this Element. 
     This break criterium checks for a maximum in the measured normal force data. 
     When a maximum has been detected the Job will jump to the next element. 
     The number of data points that is taken into account for finding (calculating) the 
     maximum can be modified by the user.

- All Rotation and Axial Measurement Elements
    - The number format for displaying the set values (for shear stress, shear rate, normal 
      force, etc.) in the JobEditor and the Job information list has been changed. 
      Small values (< 0.01) will now be displayed correctly.

- Curve Discussion Evaluation Element

- The QC part of the evaluation in this element did not work properly. This has been solved.

- JobManager and Datamanager

- /Debug command line parameter

- Both JobManager and DataManager now check if the rheowin.log file can be written in 
      the main RheoWin3 directory. If this it not possilbe, because of restricted Windows 
      user rights for example , RheoWin will try to write the rheowin.log file in the 
      RheoWin3\Drivers directory and if that is not possible it will use the Windows 
      Temporary directory.

At startup both JobManager and DataManager will now display a MessageBox 
      showing the directory where the rheowin.log file will be saved.

- DataManager

- Layout out file when loading

- The option Layout out file when loading (General page of the Preferences dialog of the
      Configuration menu) did not work when a file was loaded by double clicking on that file
      in the Windows explorer. This has been solved.
  - DataSelect for Evaluation routines

- When selecting the phase angle δ as a quantity, RheoWin selected the raw phase

angle δ-R instead. This has been solved.

- Standard page layout

- There was a problem when activating the table display after using the saved standard
      page layout when that page layout did not contain a table. 
      This has been solved.

- FileSelect for removing files

- The file select dialog for removing files from a page now allows vertical scrolling and the 
      size of the dialog can now be changed.

- Graph / Table display

- Graph display

- There was a problem in the graph display. When a certain quantity was selected for a 
      a log scaled axis and this quantity was not available in the current data set, RheoWin

might crash. This has been solved.

- The dialogs for adding text to a graph and for changing the scaling of an axis (right-
      mouse button popup window) could be opened more then once at the same time. 
      That caused confusion when the dialogs were hidden by other windows.
      Now these dialogs are "always on top" and cannot be opened more then once at the
      same time anymore.

2.13 Version 3.40

Below is a list of things that are new and/or changed in version 3.40 compared with 
version 3.30 of RheoWin.

- Installation

- Thermo Scientific

- Some bitmap images in the installation program have been modified to reflect the 
      merger between Thermo Electron Corporation and Fisher Scientific.

- Font installation

- The installation of the RheoWin font is now performed at the end of the installation 
      after all the program files have been copied, installed and registered. The advantage 
      of this is that when the font installation fails for some reason, the RheoWin installation 
      will be ok and RheoWin will run properly.

- General

- Thermo Scientific
    - The startup splash window has been modified to reflect the merger between 
      Thermo Electron Corporation and Fisher Scientific.

- Device drivers

- MARS

- Under certain circumstances problems could occur at the element transitions

between OSC Element when the AutoTension option on the Axial page was used.
      This has been solved.

- We recommend to update the MARS Firmware to the following new versions:
                    MARS Firmware version: UPD_MARS_MBOX_12_01_000.h86
                    MARS Firmware version: UPD_MARS_CBOX_12_00_000.h86

The MARS Firmware is available for download on www.rheowin.com.
  - RheoStress 600

- The MSC calibration is now performed using CR-Mode and without a sample, that

is "in air", just like with the MARS.
    - The MTC-Calibration routines from the MARS are now also available for the RS600.

Please Note: For the MTC to work a new RS600 Firmware version is needed! 
                           This version will be available in december 2006.

- RV1 driver

- The maximum allowed angular velocity can now be changed (within certain limits) using
      an *.ini file command. See the drivers.ini file in the \RheoWin3\drivers directory.
  - RheoScope Module

- This version comes with some new driver files for the camera of the RheoScope module
      which uses the Firewire (IEEE 1394) interface.

For the RheoScope Module a new MARS M-Box Firmware is needed:

MARS Firmware version: UPD_MARS_MBOX_12_01_000.h86

- JobManager

- Settings for E-Mail

- On the new E-Mail page in the Preferences dialog of the Configuration menu, the user 
      can enter the information that is needed to send an E-Mail using the Message Element 
      (see below). 
      The E-Mail setting are stored and used for each RheoWin user individually!
      Please note: Thermo does not supply any e-mail services to customers. In order to

be able to send an E-Mail using the Message Element the user has to 
                           supply a valid e-mail adress and the corresponding user name, 
                           password and SMTP server.

- JobEditor and JobController windows

- Users sometimes have a hard time finding the JobEditor window of a running Job, or the 
      JobController of a running Job when there too many windows open on the computer 
      screen. In order to solve that problem there is now a button Show JobController in the
      JobEditor (only when the Job is running!), klicking on that button will bring the 
      corresponding JobController window to the foreground. In the JobController there is now 
      a button Show JobEditor, clicking on that button will bring the corresponding JobEditor 
      to the foreground.
    - There is now also a drop-down list, that shows all currently open JobController windows,
      in the symbol bar of the JobManager, clicking on an entry in this list will put the 
      corresponding JobController window in the foreground.
  - Signatures (21 CFR Part 11 Module only)
    - After running a Job the JobEditor would "forget" the signatures in the Job and allow
      editing the Job file again. This has been solved.
      Please note: Since this is a crucial "bug" we have a fix for this bug for RheoWin 3.23 
                           and RheoWin 3.30 also.

- Manual control / Monitor Mode

- The normal force display and control are now in the left part of the dialog (and no 
      longer under the big character display).

- The big character display is now on a page with the title "Numeric" (just like the two
      graphs and the two tables) on the right hand side of the manual controls.

- Lift Element

- On the Rotation page of the Lift Element editor the user can now select to automatically
      set the absolute angular position of the motor axis to the optimal position. When this
      option is selected the lift will stop 10 mm before the measuring gap is reached, it will 
      then set the optimal angular position and automatically determine a MSC-Offset value. 
      This takes about 30 seconds. 
    - The option Find and set zero point + move lift apart has been split up in two options; 
      Find and set zero point and Move lift apart.
      When an old (existing) Job is loaded in which the original option was checked, both 
      new options will be checked automatically when loading this Job.
      Using the option Find and set zero point only, in combination with the Go to standby 
      position at in one (or two subsequent) Lift-Element(s) is usefull (with the HAAKE 
      MARS) when you do not want to move the Lift completely upwards after determining 
      the zero point. 
      Please note: Moving the lift apart to a defined position is very slow on all instruments
                           but the HAAKE MARS.

- Message Element

- The message element has been completely rewritten. 
    - The user can now define a font size and a text colour for multiple-line messages that 
      will be displayed on the screen.

- The message element can now send an E-Mail directly from a Job. This is usefull

for longer jobs that run unattended (in the lab), in this case the operator can use the
      message element to have RheoWin send him an e-mail (to his office PC) when the 
      Job is finished.

- The E-Mails are send using a very small separate program that runs completely 
      independend from RheoWin JobManager.
  - Run or Execute Element

- Optionally this element and therefore the Job can now wait for the external program to be 
      closed and/or finished.

- All measurement Elements & Temperature Set Element

- The information/status display of the "waiting for a certain temperature" part of these
      elements is now integrated in the JobController display (and not in a separate window
      anymore).

- ROT Time Element, ROT Creep Element, ROT Recovery Element

- In these Elements the user can now define an integration time for the measurement of the
      main measured values.

- ROT Temperature Step Element

- In these Elements the user can now define an integration time for the measurement of the
      main measured values just like in the ROT Step Element.

- All OSC Elements

- The maximum number of repetitions per data point is now 25 (for the MARS only). 
      The latest MARS firmware version is needed for this work.
                MARS Firmware version: UPD_MARS_CBOX_12_00_000.h86
  - OSC Freq Sweep Element, OSC Amp. Sweep Element

- The Axial page with its ThermoGap and AutoTension functionality that was already 
      available in the OSC Time Element, OSC Temp Ramp Element and the 
      OSC Temp Step Element is now also available in the Element editors of these 
      two elements.

- DataManager

- Cross-Over Evaluation

- The X-Axis range selection now works for any quantity on the X-Axis, i.e. not just for the 
      frequency or the angular frequency. 
      Currently this works in the DataManager only (i.e. not in the Cross-Over Element as part
      of a Job running in the JobManager!).
  - TTS evalution module

- The functionality of the TTS evaluation module has been expanded. On the new 
      Layout page of the TTS dialog the user can now select whether the shifted curves, 
      the shift factors and the master curve should be displayed in a new and separate

graph each. The user can (should) pre-select select a graph template for each graph.
      Predefined templates for the three graphs are part of the RheoWin installation.

- Internal Job Protocoll

- The internal Job Protocoll that RheoWin stores in all its Data files is now accessable 
      by clicking on the empty space in the Toolbar while keeping the CTRL or SHIFT key 
      on the keyboard pressed.

- Reports

- For the graphs in Reports we now use the EMF format instead of the WMF format. 
    This solves the display problems that could occur when using the exponential format 
    for the axis labels.

2.14 Version 3.50

Below is a list of things that are new and/or changed in version 3.50 compared with 
version 3.40 of RheoWin.

- Installation

- Entering the registration key information

- The user can now enter multiple registration keys during the installation.

- The maximum number of registration keys that can be entered is 20.

- Multiple rheometer and viscometer keys can be entered during the initial installation 
      as well at a later point by running the installation program again .

- Multiple optional module registration keys can be entered during the initial installation 
      as well at a later point by running the installation program again .

- When the 098-5019 key for installing all instrument is used the TTS, Spectra and
      MWD modules are installed automatically.

- New measuring geometries

- The two rotors FL26 CMC and FL29,5 CMC (for the Construction Material Cell) have 
      been added to the installation. 
      The A and M-factors of these sensors are set to 1.0, since these are so-called 
      relative measuring geometries.
    - The new Plate-Ring PR35-27 measuring geometry has been added to the installation.

- Units

- Depending on which rheometer(s) and/or which optional module are being installed 
      units that are not commonly used for the installation are now "disabled". 
      Disabled unit do not appear in the unit select dialogs for the graphs and tables. 
      Disabled units can easily be enabled in the Configuration Units dialog.

- Report templates

- All report templates are now stored in the format of the currently used Report tool
      (List&Label) version.

- Graph templates

- All graph templates will now produce graphs with a completely white background 
      (no grey rim and grey legend anymore).

- Configuration

- Units

- For the quantity torque, the unit nNm has been added.

- For the quantity stress, the unit dyn/cm² has been added.

- Rheometer drivers

- RS100 and RT10

- When running Jobs these instruments would report an "Unknown command"
      "LiftSetTension". This has been solved.

- DeviceManager

- RheoStress 1, RheoStress 300, RheoStress 600 and MARS

- The AutoStrain switch in the DeviceManager has been removed. The user can 
      now choose between CD-Mode and CD-AutoStrain-Mode in the oscillation 
      element editors.

- MARS

- The lowest allowed value for the angular speed is now n = 10^-5 rpm according to the 
      latest MARS specifications.

- The MSC Test measurement is now only performed in one rotational direction.

- The MSC Test measurement data is saved in a file with the filename Mars II-t.msc, 
      the MSC calibration measurement date is saved in a file with the filename MARS II.msc.

- JobManager

- Save-Element
    - In RheoWin 3.40, when the dialog for selecting the data file name was not set 
      to appear at the end of the job, the file names of the oscillation rawdata files 
      (*.rwr files) were "empty". This has been solved.

- Saving data files on a network sometimes caused problems. We think that this 
      problem is now solved.

- Export Element

- The setting of the checkbox Jobname was not saved. This has been solved.

- The export element now also saves to an ASCII file when there is no measurement 
      data available.
  - Temperature Set Element

- On the new Axial page of the editor of this element the user can now activate 
      the normal force control (of the MARS).

- The actual value of the normal-force signal is now displayed in the JobController 
      window when the temperature set element is active.

- Message Element

- In RheoWin 3.40, the message text was not displayed completely under certain 
      circumstances . This has been solved.

- Measurement elements icons in Job-Editor

- All elements in which a set variable is changed during the measurement, i.e. for

example the ROT-Ramp-Element, ROT-Step-Element, OSC-Frequency-Element, etc.

will now show a different icon in the Job-Editor depending on the whether the start

value is larger than the end value or vice-versa.

- Rotation Step Element and Rotation Temperatur Step Element

- These elements would crash RheoWin when in the measurement geometry used 
      in the Job the diameter was set to zero.
      This has been solved.

- Oscillation Elements

- In the oscillation elements in which the user could select between CS-Mode and 
      CD-Mode, the user can now select between CS-Mode, CD-Mode and 
      CD-AutoStrain-Mode. The AutoStrain switch in the DeviceManager has been removed.

- Oscillation and Axial Elements

- The number format for displaying the set values (for stress, strain, normal 
      force, gap, etc.) in the JobEditor and the Job information list has been changed. 
      Small values (< 0.01) will now be displayed correctly.

- CrossOver Evaluation Element

- The CrossOver Element now automatically finds upto 3 cross-over points.

- The X-range selection now works for all quantities.

- A marker line will now be plotted for any selected X-axis quantity.

- Interpolation and CurveDiscussion Elements

- The Method selection list is now larger and shows all the options at the same time.

- Some textual changes to the dialogs were made.

- DataManager

- *.RWR Viewer program

- The separate *.RWR Viewer program can now be started from the Program menu.

- TTS calculation
    - Templates for the different result graphs are now part of the software and 
      automatically installed.

- The options on the Layout page of the dialog are now active by default.

- The list of files from which the master curve was calculated is now listed in the 
      evaluation part of the Job information window.

- Spectra calculation

- The spectra calculation functionality from RheoSoft has been integrated in RheoWin.

- Compared with RheoSoft the user interface has been improved.
    - Templates for the result graphs are part of the software and automatically installed.

- MWD calculation

- The MWD calculation functionality from RheoSoft has been integrated in RheoWin.

- Compared with RheoSoft the user interface has been improved, the user can now 
      edit and save the polymer material table.
    - A template for the result graph is part of the software and automatically installed.

- Graph / Table display

- Graph display

- By right clicking on the middle part of an axis and selecting the new commands
      Exponential format or Floating point format the user can now quickly change between
      exponential format or floating point format for the axis labels.

- Newly created graphs in DataManager now are completely white, ie. there is no grey 
      rim anymore and the legend-box is white too.

- In JobManager the legend is still grey and the grey rim is still present (by default), 
      we will change those default settings in the next release.

2.15 Version 3.50.0011

Below is a list of things that are new and/or changed in version 3.50.0011 compared with 
version 3.50 of RheoWin.

- Installation

- Graph templates for TTS, Spectra and MWD

- The graph templates for the TTS, Spectra and MWD modules are now installed using 
      the complete path to these files.

- The graph template for the TTS modules have been modified. The graphs will now show
      G', G" and the phase angle as a function of the frequency.

- CR-ROT Element and the Viscostesters VT6/7L/R (plus).

- When installing RheoWin for the VT6/7 L/R or VT7plus L/R the CR-ROT ramp element
      is NOT installed anymore in order to prevent confusion since these instrument only work 
      at a limited number of rotational speeds and therefore can not perform a CR-Ramp.

- Rheometer drivers

- Drivers for RS100 (1Ncm and 5 Ncm) and RS50

- These drivers now have a build-in internal delay for the serial communication (RS232) 
      of 10 ms. This setting will NOT slow the data-acquisition with RheoWin, but might help

in avoiding communication problems. The user can always set the delay to 0 ms by 
      adding the following lines to the drivers.ini file in the RheoWin3\drivers directory.

[RS101]
              ComWait=0

The above is for the RS100 1Ncm, for the RS100 5Ncm replace [RS101] by [RS105],
      for the RS50 replace [RS101] by [RS50].

By modifying the number behind ComWait= the user can change the delay time. 
      The value is in ms, so ComWait=10, results in a 10 millisecond delay

- Drivers for RheoScope Module

- We found some problems in the third party supplied camera drivers that are used

with the RheoScope modul. This has been solved.

- Driver for MARS

- The trigger function did not work in all oscillation elements, this has been solved.

- There were problems with the Axial Ramp Element. These have been solved.

Please note: Firmware version 12.02.000 or newer is needed for the axial ramp

element to work correctly.
                            The firmware file UPD_MARS_CBOX_12_02_000.h86 is 
                            available for download at www.rheowin.com/firmware

- Drivers for RheoScope 1, RheoStres 1, 300, 600

- These drivers now send the compliance value of the currently used measuring 
      geometry to the instrument.
      Please note: Firmware version 02.12.022 or newer is needed for the compliance 
                            correction to work.
                            The firmware file Rxx_B_02_12_022_dsp_B_02_01_005.h86 is 
                            available for download at www.rheowin.com/firmware

- Driver for RV1

- An internal job can now contain 5 (instead of 3) Lift-Elements.

Please note: RV1 Firmware version 02.12.021 or newer is needed for this work!

The firmware file RV1_A_02_12_021.h86 is available for download 
                           at www.rheowin.com/firmware

- Drivers for all Viscotesters

- In RheoWin 3.40 and RheoWin 3.50 a command unknown to all Viscotesters was send 
      to these instruments causing problems in the Monitor Window. This has been solved.

- Drivers for VT7L, VT7R, VT7Lplus and VT7Rplus

- The torque offset function (option in DeviceManager) did not work. This has been solved.

- Drivers for VT7Lplus and VT7Rplus

- The spindle of these Viscostesters would start to rotate and continue to rotate after 
      using the Test or Diagnose button in DeviceManager. This has been solved.

- When running a Job with multiple CR-ROT time elements in which, for example the 1st 
      Element sets a rotational speed of 50 rpm and the 2nd Element sets a rotational speed 
      of 0.0 rpm, RheoWin will show the previously set rot. speed (50 prm) as the rot. speed

during the 2nd Element. This is due to a bug in the firmware, but this has been fixed with
      some modifications in the RheoWin driver.

- When using the VT7Lplus and VT7Rplus with RheoWin the sensor (measuring 
      geometry) MUST have the same name as the name used in the firmware of these 
      instruments!!! This means that you should NOT rename any sensor names for these 
      instruments. However you can put some characters behind the original sensor name, for 
      example “LCP (L)” or “LCP (R)” instead of “LCP” is now allowed. RheoWin will only use
      the part of the RheoWin sensor name in front the following separator characters:
      ' ' (space) , '_', (underscore) '-' (minus) und '(' (opening paranthesis).
      For example the following sensor names (for the LCP sensor) will now work: 
      "LCP (L)", "LCP_(L)", "LCP-(L)" and "LCP(L)".

- JobManager

- DataManager icon in toolbar

- The icon for starting the DataManager from the toolbar has been updated (this was still 
      the RheoWin 2.xx icon).

- Crossover Evaluation Element
    - The Available data selection "tree" on the first page of the element editor (in JobManager
      only) has been modified and now shows entries for all possible data (G", G' =f(t),

G",G'=f(f), G",G'=f(T), etc.) for easier use.

- The crossover evaluation would not work on a master curve created by the TTS module.
      This has been solved.

- Curve-Fit Evaluation Element
    - Under certain circumstances there was a problem with the Ostwald de Waele (Visc) 
      model when the Extrapolation option was active. This has been solved.

- DataManager

- JobManager icon in toolbar

- The icon for starting the JobManager from the toolbar has been updated (this was still 
      the RheoWin 2.xx icon).

- TTS, Spectra and MWD modules
    - There are now "proper" icons for these modules in DataManager.

- There was an error in the calculation of the Relaxation Spectra, the results were wrong!
      This has been fixed.

- There is a mathematical problem in the calculation of the Retardation Spectra. Because
      of that we have disabled this functionality until further notice. 
      We are working on a solution.

- TTS module
    - When the TTS calculation encountered an error RheoWin would still show (wrong) 
      results. This has been solved.

2.15 Version 3.50.0012

Compared with 3.50.0011 this version only fixes a bug in the installation. Nothing has been modified in RheoWin itself, apart from the language file

- Installation

- RheoStress 1 and RheoScope 1

- The units files were not installed when RheoWin was installed with a key for 
      these instruments.

- General

- Japanese language

There was something wrong with the japanese language in the language file.

This has been solved.

----------

3 Limitations and known issues

- RheoWin 3 can not be installed and consequently will not run on Windows 95.

- RheoWin 3 can be installed and will run on the following operating systems :

- Windows 98 with Y2K update and Internet Explorer 4.01 SP2

- Windows 98 SE

- Windows ME

- Windows NT4.0 with Service Pack 6 and Internet Explorer 6

- Windows 2000 with Service Pack 1
  - Windows XP with or without any Service Pack.
- On Windows NT4.0, Windows 2000 and Windows XP administrator priviliges are 
  needed to be able to install RheoWin 3.x.

- Please note that according to the Microsoft Windows life-cycle policy Windows 98,
  Windows 98 SE, Windows ME and Windows NT 4.0 are not supported by Microsoft 
  anymore. We therefore strongly recommend to use Windows 2000 and Windows XP 
  (see above) only.
- Some languages (in the language file) apart from German and English are 
  not 100% complete.

An updated language file will be made available on the RheoWin web-site.

- When RheoWin is installed for a RS100-5Ncm or RS50, the installation program will ask

for the diskette containing the calibration file. If you do not have a calibration file just

click Ok or Cancel to finish the installation. RheoWin is not able to perform any 
  measurement using a RS100-5Ncm or RS50 without the correct calibration file !

Please note that you can copy the calibration file into the RheoWin Drivers directory

manually after the installation was completed.

- It is strongly recommended to reboot the PC after installing RheoWin! The main reason

for this is that sometimes the RheoWin font will not be installed correctly unless the PC

is rebooted. When the RheoWin font is not installed correctly graphs and table will not

be displayed and the JobEditor and Units dialog windows will be empty!

- In the case that the installation program continues to show the message that it found 
  a previous installation although the previous installed version was already de-installed the, 
  there are still some remaining entries of an earlier previous installation in the windows 
  registry.
  In this case AND ONLY IN THIS CASE delete the following keys (and all its subkeys etc.)
  from the windows registry : 
          HKEY_LOCAL_MACHINE\\Software\\Thermo Haake\\RheoWin3
  and/or

HKEY_LOCAL_MACHINE\\Software\\Thermo Electron\\RheoWin3
  Make sure not to delete any other keys !!
  After deleting this or these keys please run the installation program again.

- RheoWin 3.21 does NOT come with drivers for the VT500, VT501 and all the RV20/30

related instruments. As a consequence error messages will pop-up in DeviceManager when

the settings of version 3.14 (or earlier) were transfered to the new installation when 
  RheoWin 3.14 was installed using the general key and the above mentioned instruments
  were installed with RheoWin 3.14. In this case one has to remove the entries for those

instruments from the DeviceManager which is a tedious job because of the appearance

of multiple error messages.

- When a Job is running in RheoWin JobManager the Windows TaskManager might show a 
  high percentage of CPU usage by the JobManager. This only means that JobManager has
  been given that percentage by the Windows Operating System because no other program 
  needs a lot of CPU power.

----------

4 Contact Thermo Fisher Scientific

If you have any questions and or suggestion regarding HAAKE RheoWin

please send an e-mail to:

support.ct.de@thermo.com

----------

This file was last updated on: 14.06.2007 by Jint Nijman